[image: image1.jpg]7000383




VANDAGRAPH  
VN202 mkII 
O2 Analyser

VM3/COP/40.70

	Date: 3 Aug 2011
	Revision Date: 8-Aug-11
	Issue: 1



Parts Used

	Pt. Number

OEM number

Description
	Qty
	Description
	
	Pt. Number

OEM number

Description
	Qty
	Description

	7930102

7000383

Enclosure Top
	1
	[image: image2.jpg]7000384





	
	7930110

8000184

Label Front Housing
	1
	[image: image3.jpg]7000385






	7930103

7000384

Enclosure bottom
	1
	[image: image4.jpg]CALIBRATION

ON/OFF

8000184

7930108




	
	7930112

8000186

Label Battery Housing front
	1
	[image: image5.jpg]8000185





	7930104

7000385

Enclosure battery cover
	1
	[image: image6.jpg]VANDAGRAPH LTD.

15 Station Road, Cross Hills,

4. Kcighley, W. Yorkshire,

BD20 7DT, U. K
Tel : +44 (0) 1535 634900
Fax: +44 (0) 1535 635582

www. vandagraph. co. uk

o X ce

8000186




	
	7930111

8000185

Label Back Housing
	1
	[image: image7.jpg]9000052

I

MR





	Serial number Label
	
	[image: image8.jpg]9000052

7930119





	
	9730040

P16NP10KMA

Pot 10K Cal control
	1
	[image: image9.jpg]7931022

7000386





	9000052-1

PCB loaded
	1
	[image: image10.jpg]\ g

7930128

8000042




	
	9730036

150-243

Yellow Push Switch, Round,
	1
	[image: image11.jpg]i

7000248





	7930119

9000052

Graphics Display
	1
	[image: image12.jpg]5000244

7930125




	
	9730037

148-921

Switch waterproof kit
	1
	[image: image13.jpg]0000000





	7930122

7000386

Holder Display 
	1
	[image: image14.jpg]7000085




	
	7930128

8000042

12 x12 rubber cushion
	1
	[image: image15.jpg]7930108 VE
7 ". g /B

7000086





	7930114

7000248

2.0 x h6.0 x3.5 screws PCB holder
	4
	[image: image16.jpg]7000390




	
	7930125

5000244

9 Pin wire 
	1
	[image: image17.jpg]5000243





	9080015

121-0467

Heat shrink Tubing

 1 cm x 2mm
	3
	[image: image18.jpg]7930113

7000247



 


	
	7930127

5000245

Electronic wire 

Any colour
	1
	[image: image19.jpg]VN202 mkll

% OXYGEN

7000389





	7930107

7000085

CO2 +v Battery term
	2
	[image: image20.jpg]7000375




	
	7930113

7000247

1.7 v h4.0  3.0 screws pcb
	2
	[image: image21.jpg]7930123

9000053






	7930108

7000086

CO2  -v Battery term
	2
	[image: image22.jpg]9000054




	
	7930124

9000054

PCBA JTAG-B-V1.1
	1
	[image: image23.jpg]crew:7000375 ¢2.6Xh10.0

crew:7000376 ¢2.6Xh6.0






	7930105

7000390

Seal ring Housing
	
	[image: image24.jpg]



	
	7930126

5000244

8 pin wire
	1
	[image: image25.jpg]




	7930116

7000376

2.6 x h6.0  4.0 screws housing
	
	[image: image26.jpg]



	
	7000389

Display window
	
	[image: image27.jpg]BATTERY TYPE 1.5V AA ALKALINE
VN202 mkll

Pt. No. 7910100

SN 00003






	7930123

9000053

PCBA JTAG-A-V1.1
	
	[image: image28.jpg]



	
	793115

7000375

2.6 x h10.0  4 screws housing
	4
	[image: image29.jpg]




	
	
	
	
	Double sided taps
	1
	

	
	
	
	
	
	
	


1.0 Unwrap the cases and separate them into their component parts, 

2.0 Add the Main label to the front panel

3.0 Fit the yellow switch 9730036 for auto switch off or( Green for non auto switch off). 
Ensure the waterproof kit 9730037 is fitted including the O ring

4.0 Add the Calibration potentiometer 9730040
Ensure the long centre terminal is at the top.
The white spot should also be at the top half way between max and min settings.

5.0 Graphics Display
If the graphics display and PCB are separate fix the graphics display to the back of the PCB with double-sided tape

Check OEM as pcb may need shorting links removed. See end of procedure


READ 6.0 > 15.0 before proceeding

6.0 Remove the LCD protective film or enough to allow it to be removed through the front panel hole

7.0 Fit the 9 pin cable 7930125 to the 9 pin socket.

8.0 [image: image30.jpg]


Locate the PCB/LCD assembly in the recess as near to the 

9.0 Add the frame 7931022. It only locates one way.

10.0 Place the small foam pad underneath the frame 

11.0 Fasten with 4 off 7930144 silver screws. 

12.0 NB Do not tighten as the LCD may need realigning.
 Check once the wiring is completed and the batteries can be fitted.

13.0 Blue and Red go down the side of the enclosure 

14.0 Grey & Purple go under the frame to the right

15.0 The Yellow, Orange, & Black go under the frame to the left

16.0 [image: image31.jpg]


Green and Brown go through the small channel in the Jack enclosure

17.0 Wire Black & Grey to the switch

18.0 Wire Orange Yellow Purple to the Calibration pot.

19.0 Add heat shrink tubing 1cm x 2mm dia

20.0 Wire the Jack socket. The short White & Brown go through the groove and into the Jack socket compartment.

21.0 Brown (may be Green) pin 4 on socket    White Pin 5 on socket

22.0 Fit the Jack socket Type 42 into the small area and screw on the nut

23.0 [image: image32.jpg]GD
B
X oyEm

:l

H
j"E%EEa 5
l Eaxa‘ﬂ“mm

mjuzi.hn Le“_l
| & | _l—l_ E]r‘“"
GV!I"DI 5§ IU‘-’




Ensure the wires go through the access slit wires do not get 
trapped and are tucked down around the socket so that they do not foul the jack plug when it is inserted

24.0 Connect the Red to +ve and the Blue to -ve on the battery terminals

25.0 Connect a wire across the lower battery terminals

26.0 Add non slump silcone to the access holes of battery cover 

27.0 [image: image33.jpg]


CHECK
The seal is in the groove all around the enclosure
The Battery wires are not going to be trapped
The wires to the Jack are in their groove and covered by the seal.
The LCD is aligned
The potentiometer increases the reading when turned clockwise. A sensor or Oxycal is required for this test

28.0 Fix the two parts of the enclosure together.

29.0 4 screws 793115 are used in the top 
half of 
the enclosure and 2 screws 7930116 in the bottom two holes.

30.0 Fix the Rear Label

31.0 Front small label

32.0 Check Software 

Software is different for different OEM

33.0 Add serial number label

34.0 Fit the battery cover 40.


R11=10k      R12= NC no auto off  }  NC means cut out the resistor

R11=NC      R12=0R  Auto off       }  Leave resistor in place it is zero ohms 

R13=10K    R14 = NC    reverse display I.e.  reads 0% when it is at 100% for non oxygen gases

R13= NC    R14= 0R      normal display


VN202 mkII manufacture.odt    8/8/2011        Page 4 of 4

