Parts Used
Pt. Number / OEM number / Description	Qty	Description		Pt. Number / OEM number / Description	Qty	Description
7930102 / 7000383 / Enclosure Top	1			7930110 / 8000184 / Label Front Housing	1
7930103 / 7000384 / Enclosure bottom	1			7930112 / 8000186 / Label Battery Housing front	1
7930104 / 7000385 / Enclosure battery cover	1			7930111 / 8000185 / Label Back Housing	1
Serial number Label				9730040 / P16NP10KMA / Pot 10K Cal control	1
9000052-1 / PCB loaded	1			9730036 / 150-243 / Yellow Push Switch, Round,	1
7930119 / 9000052 / Graphics Display	1			9730037 / 148-921 / Switch waterproof kit	1
7930122 / 7000386 / Holder Display	1			7930128 / 8000042 / 12 x12 rubber cushion	1
7930114 / 7000248 / 2.0 x h6.0 x3.5 screws PCB holder	4			7930125 / 5000244 / 9 Pin wire	1
9080015 / 121-0467 / Heat shrink Tubing / 1 cm x 2mm	3			7930127 / 5000245 / Electronic wire / Any colour	1
7930107 / 7000085 / CO2 +v Battery term	2			7930113 / 7000247 / 1.7 v h4.0 3.0 screws pcb	2
7930108 / 7000086 / CO2 -v Battery term	2			7930124 / 9000054 / PCBA JTAG-B-V1.1	1
7930105 / 7000390 / Seal ring Housing				7930126 / 5000244 / 8 pin wire	1
7930116 / 7000376 / 2.6 x h6.0 4.0 screws housing				7000389 / Display window
7930123 / 9000053 / PCBA JTAG-A-V1.1				793115 / 7000375 / 2.6 x h10.0 4 screws housing	4
				Double sided taps	1

1. Unwrap the cases and separate them into their component parts,
2. Add the Main label to the front panel
3. Fit the yellow switch 9730036 for auto switch off or( Green for non auto switch off).
   Ensure the waterproof kit 9730037 is fitted including the O ring
4. Add the Calibration potentiometer 9730040
   Ensure the long centre terminal is at the top.
   The white spot should also be at the top half way between max and min settings.
5. Graphics Display
   If the graphics display and PCB are separate fix the graphics display to the back of the PCB with double-sided tape
   Check OEM as pcb may need shorting links removed. See end of procedure
   READ 6.0 > 15.0 before proceeding
6. Remove the LCD protective film or enough to allow it to be removed through the front panel hole
7. Fit the 9 pin cable 7930125 to the 9 pin socket.
8. Locate the PCB/LCD assembly in the recess as near to the
9. Add the frame 7931022. It only locates one way.
10. Place the small foam pad underneath the frame
11. Fasten with 4 off 7930144 silver screws.
12. NB Do not tighten as the LCD may need realigning.
    Check once the wiring is completed and the batteries can be fitted.
13. Blue and Red go down the side of the enclosure
14. Grey & Purple go under the frame to the right
15. The Yellow, Orange, & Black go under the frame to the left
16. Green and Brown go through the small channel in the Jack enclosure
17. Wire Black & Grey to the switch
18. Wire Orange Yellow Purple to the Calibration pot.
19. Add heat shrink tubing 1cm x 2mm dia
20. Wire the Jack socket. The short White & Brown go through the groove and into the Jack socket compartment.
21. Brown (may be Green) pin 4 on socket White Pin 5 on socket
22. Fit the Jack socket Type 42 into the small area and screw on the nut
23. Ensure the wires go through the access slit wires do not get	trapped and are tucked down around the socket so that they do not foul the jack plug when it is inserted
24. Connect the Red to +ve and the Blue to -ve on the battery terminals
25. Connect a wire across the lower battery terminals
26. Add non slump silcone to the access holes of battery cover
27. CHECK
    The seal is in the groove all around the enclosure
    The Battery wires are not going to be trapped
    The wires to the Jack are in their groove and covered by the seal.
    The LCD is aligned
    The potentiometer increases the reading when turned clockwise. A sensor or Oxycal is required for this test
28. Fix the two parts of the enclosure together.
29. 4 screws 793115 are used in the top	half of	the enclosure and 2 screws 7930116 in the bottom two holes.
30. Fix the Rear Label
31. Front small label
32. Check Software
    Software is different for different OEM
33. Add serial number label
34. Fit the battery cover
R11=10k R12= NC no auto off } NC means cut out the resistor
R11=NC R12=0R Auto off } Leave resistor in place it is zero ohms

R13=10K R14 = NC reverse display I.e. reads 0% when it is at 100% for non oxygen gases
R13= NC R14= 0R normal display