Before you ship it back.
We have very very few real problems with the Mixcheck. 95% involve calibration.
Can you do the following before you send it back.
1. Feed in Dry AIR
2. If the Helium reads 00.0 turn Zero control until it reads 00.2
3. Oxygen should be set to 20.9%
4. Feed in 100% helium
5. Set the span to 100% If reads less than 100% turn the span anticlockwise and if it reads over 100% turn span clockwise.
6 Oxygen should read less than 1% + It can take a long time to reach this level
7. Feed in dry Air
Oxygen should return to 21% and Helium to 00.2
Make sure you use the 2ltr flow control

I find this solves most of the problems and it should save you the postage.
However if you still have problems take some photograph showing the above sequence of steps and send them to me.

John S. Lamb
18/08/11