	0111211 QA Procedure


Model number:  GB-300 (fixed field)
S/N: (variable field)
Manufactured date: i.e. 2010 (variable field)
Part number: C77968A (fixed field)
Revision number:  i.e. 3 (variable field)
Software version: i.e. V1.1 (variable field)
Battery compartment label reflects battery door label: Y/N
Carrying case label reflects the battery door label:  Y/N

NOTE: Apart from revsion number is Rev: 4
Screen protector fitted: Y/N

Sensor output connector cover (silicone plug):  Y/N

Text above the display states PERCENT OXYGEN   Y/N

1 x GB300 analyser (yellow). Note: no cradle, no v-mount plate Y/N
1 x I-17 oxygen sensor (0110517) Y/N
1 x T piece (0120100) Y/N
1 x Flow diverter (0120099) Y/N
1 x  Connector 22mm I.D./30mm O.D. - tapered (0120133) Y/N
1 x Pack of batteries, (3 x AA) Y/N
a) Test the sensor, ensure between 7 – 13 mV (use DVM)

b) Install test batteries.

c) Upon switch on display should momentarily display 

188.8

Battery icon – all segments lit

Then display should read

00.0

Battery icon – all segments lit

d) Press on/off button, unit should switch off.

Page 1 of 2

e) Connect sensor cable, ensure secured with threaded lock.

Attach sensor to the sensor cable, ensure secured with threaded lock.

Switch unit on.

Press and hold the on/off button until the display starts to count down,

 i.e.  3.  2.  1.  0.

Then the unit should read 20.9% 
Switch off the unit and reverse the cable, then repeat the test e).

QA passed: Y/N

Add from stock.

1 x 0131202  Sensor cable Y/N
1 x 0120131  Connector 15mm I.D./22mm O.D. – tapered. Y/N

v1.1

12/09/11

Page 2 of 2
