	     [image: image2.jpg]ifm

P/N C73880 REV'-

r e S

W

_HEOEERRRERERERREREEE
FREDRRRERRNRRABARARS

l‘Illl‘l““ll“l“llllllllll‘
1 llllHllH'lllHHlllIHHHll

ANALYTICAL INSTRUHE

“
P





	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire  BD20 7DT - United Kingdom
Tel: 01535 634542               Fax: 01535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



MX300 firmware V1.4 upgrade

Summary procedure for replacing the PCB assembly

from a MX300 Oxygen Monitor.
[image: image1.png]



Part required:

1 x 
0100160
MX300-I PCB assembly,
Output: 0 - 1 V
Display: 3 1/2 digits
Assembly model: C73884-A
Software: V1.4

Tools required:

Anti-static workstation

Philips Size 0 screwdriver.

At an anti-static workstation:

1) Remove the battery cover and the three batteries.

2) Remove the ‘V’ mount adapter.

3) Remove the four screws from within the battery compartment and the single long screw located at the top of the rear panel.

4) Carefully remove the rear panel.

5) The existing PCB can then be removed, leaving the keypad membrane in position within the front panel assembly.

6) Remove the new PCB assembly from its anti static bag and remove the clear protective tape from the LCD screen prior to fitting the PCB assembly into the front panel. Ensure that the sensor input and output data connectors are correctly located.

     Insert the old PCB assembly into the antic-static bag.

7) Refit the rear panel and then refit the rear panel screws, ‘V’ mount adapter, batteries and battery cover.

8) Carry out functional test, testing all buttons and functions. Carry out calibration procedure in air and/or 100% oxygen. Note: ensure that all four backlight LEDs are functioning.

August 2008
[image: image2.jpg]