	     [image: image5.jpg]‘assnanssnas ’

o) Ehes o LN RN YN J
AT .W.mh.‘ £i5 a3n B iy





	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire  BD20 7DT - United Kingdom
Tel: 01535 634542               Fax: 01535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



AX300 Oxygen Analyser & MX300 Oxygen Monitor
Summary procedure for upgrading the casing &

oxygen sensor input connector.
Part required:

Either 

1 x 0121216 AX300 Case Upgrade Kit Rev.2

or

1 x 0121217 MX300 Case Upgrade Kit Rev.2
Each kit comprises of:

1 x Front panel Rev. 2, AX300 = 0100174,  MX300 = 0100175
1 x Rear panel Rev. 2, 0121215
[image: image1.png]


1 x Oxygen sensor input connector Rev.2, 0131214
Tools required:

Anti-static workstation.

Philips Size 0 screwdriver.

Soldering iron.

At an anti-static workstation:

1) Remove the battery cover and the three batteries.

2) Slide off and remove the ‘V’ mount adapter.

3) Remove the four screws from within the battery compartment and the single long screw located at the top of the rear panel.

4) Carefully remove the rear panel; the PCB assembly can then be removed. 

NOTE: Put the battery lock aside – to be used with the Rev.2 rear panel.

5) Unsolder the red and black sensor input wires from the existing connector. Then solder the wires to the Rev. 2 connector. For connections see the images in step 7) 

6) Transfer the keypad membrane to the Rev. 2 front panel.

P.T.O

7) Position the PCB on top of the keypad membrane and fasten the oxygen sensor input connector to the front casing. Reposition the output connector.

[image: image2.jpg](@]

PIN "criesf) R

s ;,g;a;g;;;:m:;;.'?‘

;;z,,

it

i





[image: image3.jpg]‘dddsd
igede

RPTTTIILy

L

+ $84eLantttigattety
0 v % m

5 /

"
/

(.

-

L7
o ﬂ

&

fa



[image: image4.jpg]Sitrorii
ST7063C-QG
AD4G16.1

NALYTIGA
Assy C7388%-\





8) Fit the Rev. 2 rear panel together with the existing battery door lock. Then refit the rear panel screws, ‘V’ mount adapter, batteries and battery cover.

9) 
Carry out functional test, testing all buttons and functions. Carry out calibration 

procedure in air and/or 100% oxygen.
October 2013
[image: image5.jpg]