[image: image1.jpg]


VANDAGRAPH



R22ADD Assembly

VM3/COP/40.15  R22ADD

	Date: 
	Revision Date: 25-Jan-12
	Issue: 1




Used in Dive Dynamics Oxygen Monitor    Pt Number 9212-0

Method.

1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Test Sensor

4. Remove the O2 sensor specification leaflet and discard.  

5. Remove excess outer case unto level of PCB. VM3COP40.11 Remove the Molex and connect wires with remote 2 way molex.

6. Pay attention to wire positions on the PCB Black to BLK: Red to RED

7. The lock position on the molex

8. The center is left unconnected

9. Red to Right flat side of molex

10. Label sensor with R-22ADD

11. Test for output 9-13mV.

12. Package in gas barrier bag

Parts Required�
�
�
Quantity�
Description�
Part No.�
�
�
1�
R-22MED: R-22A: R-22BUD are all equivalent)  �
R22  Sensor;�
�
�
1�
White Box & Lid�
�
�
�
1�
wire 10 cm  Black�
�
�
�
1�
wire 10 cm  Red�
�
�
�
2�
Pins  Farnell �
143-116�
�
�
1�
Shell Farnell  �
143-126  2 pin�
�
�
1�
Packing Material�
�
�
�
1�
Round “Vandagraph” Label�
�
�
�
2�
Serial number Labels�
�
�
�
1�
Polythene Bag (outer)�
�
�
�
 








R-22ADD.doc          Page 1 of 1

