VIAMED


VM2500-M Contents & Testing

VM3/COP/54.00   4410500

	Date:22/01/10
	Revision Date:31-Jan-12
	Issue: 1



VM2500-M  4410500 Contents & testing:

Contents on Receipt

1 x VM2500-M   

(Blue casing) fitted with silicone cover

Document the serial number and the software version and type, i.e. 

VM-2500-M, Version number, Europe, CD Type/Version

Check front panel is 

VM-2500-M, VIAMED

Check serial label is Viamed version

1 x IRMA CO2 analyser P/N: 4420505 

Need to document the serial number and check that ref. is 200104 – Needs bar-coding so it can be linked on QA

1 x IRMA airway adapter (adult/paediatric). 

Check that Phase-In version and not Phase-In/ Smiths BCI. Our part number for this 4420512. If Phase-In/Smiths BCI replace with our version from stock. Single Item p/n:  4420512  (Multi Pack: 4420511)

1 x USB data cable.

1 x Power supply 

Type FW7660M/06

1 x Power supply plug adapter (EU)

1 x Power supply plug adapter (UK)

1 x Li-poly battery (CT-2500) Check Battery has Viamed Version Label

4 x AA batteries

1 x CD with user manual… - Check version of BP CD, should still say viamed on

Document the CD version

1 x Quick start guide. – In house Printed

Batteries should arrive with 90% charge – Double check before QA and charge for 2-4 hours if low charge

Test as per QA Form

Software Version – I.e. V5.9  (Drop Down Medical / Veterinary) – Check that it is VM and not capnotrue Firmware 

Software Region – Generally Europe

Red Screw Protector – Y/N

IRMA CO2 Analyser Included – Y/N

Irma CO2 Analyser Barcode – Should be able to scan

IRMA Airway Adapter included – Y/N

USB Cable included – Y/N

Power Supply included – Y/N

Power Supply Plug (EU) included – Y/N

Power Supply Plug (UK) included – Y/N

Li-poly Battery  (C-2500) With Viamed Label – Y/N

4 x AA Batteries Included – Y/N

CD Version – I.e. 1.6 – This needs to be checked against the compatibility sheet – Older Firmware will not work with new PC software and Vice versa – Compatability ref is mainly for Older units that have been upgraded or downgraded. New units should arrive with current CD version

Quick start guide included – Y/N

Test with test sensor and test Irma – Y/N – Test needs to be done with AA batteries Li-ion Battery and mains – No test Irma – use units own Irma, adapter in box

Check time Date – Y/N – Double check time actually changes on main screen– Known bug with time not changing 

Delete Test data – Y/N – Very important – New section on Mast indemnity for Scotland (Not sure on others) That any device that stores data has to be clear

Packaging sleeve has large barcode and colour dot – Y/N – Found in the sticker Drawer by the testing bench

Pass – Y/N

Additional things to test –

Power unit on and wait 10 seconds before attaching the Finger probe

Disconnect the finger probe while the unit is powered on and wait 10 seconds before connecting – Unit should detect probe and pick up signal again

Check the display Brightness and compare to another unit, on the same level of brightness

Check the sound of the unit – On power on, Pulse Tone and alarms – Should be a clear sound

CO2 unit settings and alarm limits 

Medical Units – leave as they are – CO2 unit needs to be confirmed on a per order basis

Veterinary Units – Set default Adult alarm for BPM to 250 and set CO2 unit to mmHg

Page 2 of 2

