[image: image3.png]



[image: image1.png]QM .2000 ~ 6;}

SAN=-0Q-SUB [TALIA s..1 &
SINCERT

EQUIPMENT FOR GASES K
2

STABILIMENTO ED AMMINISTRAZIONE:
via L. da Vinci, 168 - 20090 Trezzano sul Naviglio (Mi)
Tel. 02.445.20.75 - 02.445.40.57 - Fax 02.445.06.34 - www.sanosub.com





ATATTENTION : IT’S EXTREMELY IMPORTANT TO READ AND TO UNDERSTAND ALLA THE   INFORMATIONS CONTAINED IN THIS MANUAL IN ORDER TO EXTEND THE KNOWLEDGE OF THE APPARATUS TO AVOID EXPLOSIONS OR SERIUOS INJURES  

[image: image2.wmf] 0123, directive 93/42 CE 
SAN-O-SUB ITALIA                          OXYGEN REGULATOR FOR MEDICAL GASES

	COD.
	MODEL
	CONNECTION  IN/OUT
	PRESSURE

GAUGE IN/OUT
	DIMENSIONS

	RO.200
	Preset pressure reducer with fixed outlet 1 .manometer
	UNI (Any European standard / 9/16”
	200 /  4        315-bar
	90x90x130

mm.

	RO.200
	Adjustable pressure reducer with  2 manometres
	UNI (Any European standard / 9/16
	200/  4         315- 16bar
	90x90x130

mm

	RO.300
	Adjustable pressure reducer with index flowmeter
	UNI(Any European standard / 9/16
	200/  15 lt./ min  315bar
	90x90x130

mm

	RO.200
	Adjustable pressure reducer with flowmeter
	UNI(Any European standard / 9/16
	200/  15 lt./min   315bar

200/  06 lt./min   315

bar
	160x170x130

mm


Accuracy: + or – 10% o 0,5lt.min.

Pressure release valve factory preset. 

Outlet of regulator: 3.5 mm hose or with quick connection AFNOR/UNI…

Class of manometers: 2.5

DESCRIPTION

Pressure reducer model RO.200/300 are manufactured in brass EN 737/1 or in aluminium alloy, chromed externally and assembled carefully according to the model. All the models have pressure relief valve preset from the manufacturer and cannot be tampered with. This is to be used in the control and distribution of medical gases.

The reducer is suitable for assembly on medical cylinders in accordance with existing standards. UNI/DIN/EN/NF or PIN INDEX and with all the norms that rule gas distribution.

Before use check that the reducer is suitable for the same medical gas indicated on the cylinder.

The reducer is to be used with max pressure of 232 bar and in case of preset pressure, outlet is 4 bar.

IDENTIFICATION

On the body of pressure reducer there are the following markings:

-
Name or symbol of gas

-
Model with month, year, series number

-
CE-with number of notified organisations

-
Name or initial of producer

-
Maximum inlet pressure

USE

Before connecting the reducer to the cylinder follow these procedures:

1.
Check that the reducer is compatible with same gas of cylinder.

2.
Fasten the cylinder on the wall or use an appropriate trolley.

3.
Check that the seal is present on the medical valve.

4.
Remove the seal and check that the thread is not damaged.

5.
There should be no traces of OIL, GREASE or DIRT on the reducer. 

ATTENTION: In case there is the presence of any one of the above substances, don’t 

use pressure reducer. It is extremely dangerous and may cause injury. Check that there 

is O-ring on the connection and the gauges are in good condition. In the case of dirty 

or damaged O-ring, don’t use regulator. Replace the O-ring following the instructions 

under MAINTENANCE.

6. Screw the swivel connector into the cylinder valve with a force of 10/15 Nm to assure 

enough seal. If you note a blocking during this operation don’t force with a key because probably one of the two threads is damaged.

Fasten the reducer with h.p.manometer in upright position.

7.
Unscrew handwheel pressure regulator (in case of adjustable pressure regulator).

8.
Connect at the outlet of regulator the medical device that you require by means of rubber tubing, checking that they support the outlet pressure of regulator and that the medical gas is compatible.

9.
Screw handwheel of flowmeter in case of reducer with flowmeter.

10.
Stay on side of regulator and open slowly the cylinder valve one complete turn. 

ATTENTION:  opening quickly the valve may cause autoignition of the internal 

components.

11.
Turn anticlockwise the handwheel of flowmeter and clockwise the handle of regulator 

to start the gas supply. 

In case of regulator with flow index the flow must be read on the index. 

In case of regulator with flowmeter the flow must be read by seeing that the middle of 

the ball of the flow indicator is on the value required.

In case of regulator with low pressure gauge the outlet pressure must be read on the 

manometer.

12.
Check that there is no leakage from the pressure relief valve. In case of leakage stop 

the use of regulator, close the cylinder valve and return the pressure reducer for 

maintenance.

13.
Checking of eventual leakages from pressure reducer:

Close pressure reducer in one of the following ways:

a) Turn the handwheel of flowmeter clockwise, the flow ball indicator starts to fall 

down to zero (in case of pressure reducer with flowmeter).

b) Turn the handwheel of index flowmeter anticlockwise (in case of pressure reducer 

with index flowmeter).

c) Turn the handwheel of pressure reducer anticlockwise (in case of adjustable 

pressure reducer with 2 manometers).

d) Close flow supply (in case of pressure reducer only). Close cylinder valve and wait 60 sec.

High pressure manometer must remain at same level of working pressure. In case 

there is a considerable decrease, the fault is a leakage in the unit.

14.
Attention: in case of N2O or CO2 medical gas being used for more than 60 min. there 

may be a freezing of the unit that is seen as frost on the reducer. In this case close the 

cylinder valve and wait until unit returns to normal condition.

AFTER THE USE

Close cylinder valve and open the reducer to purge residual gas before unscrewing the reducer from the valve. ATTENTION: be sure there is no pressure left in the reducer before unscrewing it from the valve. If it is difficult to unscrew the reducer, it is possible that there is still pressure inside.

PRECAUTIONS FOR USE WITH OXYGEN

INCORRECT USE MAY RESULT IN SERIOUS INJURY OR DEATH
A.
Pay particular attention when using medical oxygen gas min. 99,5% and with a sealed 

cylinder valve .Max working pressure 200 bar. 

B.
Fix oxygen cylinder to a wall or on a trolley.

C.
If you note any abnormality during use such as leakage, abnormal noise, loss of oxygen from relief valve, freezing, or faulty gauges etc. suspend use and ask for assistance from your local agent.

D.
Verify the integrity of the unit, the valve and washer of valve connection. For more 

Safety, substitute after every use.

DO NOT use near heating source or bare flame. 

DO NOT use oil.

USE pressure regulator with care. DO NOT contaminate with dust or greasy hands. 

CHECK
Verify at least every 6 months or at intervals as agreed by user (hospital or health company).

1.
Check leakage on outside of pressure reducer following point 13 abovementioned.

2.
Check all threads for leakages using a solution of water and soap.

3.
Check the accuracy of measure of flow. In this case it is best to seek technical help 

from your local agent.

MAINETENANCE

SAN-O-SUB pressure regulators are manufactured with material that does not require ordinary maintenance if used correctly and the checking every 6months is effected. Nevertheless, inside the instruments there are non-metallic parts that require an extraordinary maintenance after and every 5 years from production.

Ordinary maintenance carried out by professional technicians. ATTENTION: effect the following only with reducer removed from the cylinder.

A.
Replacement of O-Ring at inlet of pressure reducer - valve connection (use only viton 

O-Ring 70 sh.)

B.
Replacement of H.P. and L.P manometers (use only oxygen manometer diam. 40mm 

1/8”gas conical class 2,5)

Attention: When removing faulty manometers take care to clean the thread from 

residue of teflon or seal paste by using decontaminated air or nitrogen.

DUST or BRASS RESIDUE left inside pressure reducer may cause serious injury. 

C.
Clean pressure reducer with cotton cloth. 

EXTRAORDINARY maintenance must be effected after 5 years from production or in case of abnormal function resulting from periodical checks, maximum service life for  the product is 10 years providing the above maintenance recommandations are followed

This is only to be effected by SAN-O-SUB ITALIA or authorized local agent.TRACEABILITY OF REDUCER: SAN-O-SUB guarantees the traceability of reducers sold under the CE norm for medical devices.

SAN-O-SUB ITALIA Stabilimento e AmministrazioneVia L. Da Vinvi 168 TREZZANO SUL NAVIGLIO MILANO ITALY

