	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	VN202 mkII Tek-Ox Auto to Non auto conversion

	VM3/COP/42.06

	Date: 30-Sep-03
	Revision Date: 23-Feb-12
	Issue 1

	Only to be attempted by trained personel



	Parts Required

	Quantity
	Description
	Part No.
	

	1
	VN202 mkII PCB  
	9730121

may change
	

	1
	Switch Green  
	
	


Procedure 

1. PCB Conversion, 

1. [image: image1.jpg]No Auto Off

R11 = 10K
R12 = NC.

Auto Off

R11 =NC.
R12=0R




[image: image2.jpg]


Open the case by removing the six screws. NB the two at the end of the battery compartment are smaller.

2. If modifying a new instrument see section 4 for dismantling instructions 

3. [image: image3.jpg]


Modify pcb as above NB SMT components 






4. Dismantling Instrument
Remove Software pcb.
Remove Frame screws and Frame
Pcb can then be unplugged
Unsolder and remove switch
Replace with Green capped switch
Reassemble and test Unit should not switch off

VM3Cop42.06 Tek-ox VN202 mkII Non Auto to Auto conversion.doc       23/02/2012      Page 2 of 2

