VIAMED 

Oxygen Sensors


Incoming Sensor Returns Testing

VM3/COP/40.02

	Date:5-Nov-03
	Revision Date:5-Nov-03
	Issue: 1



	Procedure


1.0 Initially, read carefully any information supplied by the user.

2.0 If it is a telephone problem, then listen carefully, ask exact questions as to how the sensor is being used; whether;

a) VN202

b) Other Analyser

c) Re-breathers

d) Mixing panels

3.0 Refer to the Reference book “Oxygen Measurement for Divers” for information on:


Sensor Linearity - Chapter 10

Check “20.9%”

Check “100%” - Read chapter 20 - Calibration

Blowing across the sensor uses the CO2 to generate a low PPO2 to pull the Oxygen out of the Sensor. Ensure the reading is “20.9%”


Temperature – Chapter 11

Diving cylinders use 200 Bar pressure and the gas released will be at very low temperature (page 58).

 
Accuracy – Chapter 13, and Sensor Outputs – Chapter 15


Response Time.

4.0 Basic checks required:

Visual: - Leakage, Mechanical damage, Membrane damage

Linearity checks

End of life test: - Leave in “100%” O2 for several minutes.

Stability Test: - Leave in air overnight.










Page 1 of 1

