	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Helium  / O2 Analyser. Mixcheck

	VM3/COP/54.01

	Date: 25-Apr-03
	Revision Date: 26-Oct-06
	Issue 1

	Helium / O2 Analyser: O2 Assembly, Test and QA Procedure.



The Mixchek Helium / Oxygen Gas Analyser arrives without the O2 sensor or batteries fitted. 

	Equipment required.
	Sockets and ratchet driver, cross point screwdriver, flat blade screwdriver, kitchen tissue, Isopropyl alcohol, silicon polish.


1. Ensure there is no physical damage to the unit or marks to labels.

2. Unscrew the 2 fasteners and remove the battery cover.

	[image: image13.jpg]



	[image: image1.jpg]0 Teledyne
Helium/Oxygen Analyzer







3. Remove 4 x nuts / mylar washers and remove the instrument panel and battery box.

	[image: image2.jpg]0 Teledyne
Helium/Oxygen Analyzer





	[image: image3.jpg]





4. Place the instrument panel face down and screw the O2 sensor into the housing – finger tight.  Plug the free connector into the circuit board – the plug is keyed and only fits one way.

5. Record the O2 sensor serial number. Refit the instrument into its box  - ensure that the sample tube and battery lead are not trapped between the instrument and the box. Refit the 4 x retaining nuts / mylar washers. Fit 4 x batteries (size C, alkaline) in the correct polarity.

	[image: image4.jpg]



	[image: image5.jpg]





	[image: image6.jpg]Turn the ¥

et 02 span prior 10
uld 10

flo

holder with

10 20.9% Wit e

Attach the Sam

M, o sLPM. The 0N o
\he readings 1





	[image: image7.jpg]e\ed\J
o*‘lg

l't







6. Insert the battery box back into the bottom of the box and refit the battery cover and 2 x battery cover retaining screws / shake-proof washers.
	[image: image8.jpg](6-10VDC @ 50MA MAX )

BATTERIES INSIDE

REQUIRES 4 SIZE © ALKALINE
ANSI—14A, IEC~LR14





	[image: image9.jpg]Percent Helium

LOW BATTERY

40 Teledyne

Helium/Oxygen Analyzer

(6-10V0C @
BATTERIES IN:

REQUIRE:






	
	
	


7. Switch on the Helium / Oxygen Analyser and select Battery Test.  Ensure the analyser shows greater than 110 hrs of battery life remaining on the display.

8. Run 100% oxygen through the sample tube at 2 LPM.  Ensure that the gas runs for a sufficient length of time to purge the analyser of all traces of other gases.  Calibrate the analyser to read 100.0% oxygen by adjusting the O2 Span control.

10.  Run air through the sample tube at 2 LPM. Ensure that the gas runs for a sufficient length of time to purge the analyser of all traces of other gases.  Ensure the analyser reads 21.0±1.0% oxygen and calibrate the analyser to read 00.1% helium by adjusting the Helium Zero control.

12.  Run 100% helium through the sample tube at 2 LPM.  Ensure that the gas runs for a sufficient length of time to purge the analyser of all traces of other gases.  Calibrate the analyser to read 100.0% helium by adjusting the Helium Span control. 

13.  Run helium / Oxygen mix, through the sample tube at 2 LPM. Ensure that the gas runs for a sufficient length of time to purge the analyser of all traces of other gases. Ensure the analyser shows readings of helium / oxygen ( 1% of that from the cylinder.

	[image: image10.jpg]Percent Helium

LOW BATTERY

2\ Teledyne

Helium/Oxygen Analyzer

e S





	Gas mix used: 

27% oxygen. 32% helium.


14.  Run air through the sample tube at 2 LPM. Ensure that the gas runs for a sufficient length of time to purge the analyser of all traces of other gases. Calibrate the analyser to read 20.9% oxygen by adjusting the O2 Span control. Switch the analyser off.

15. Attach “Vandagraph” labels to the lid of the analyser box and to the front panel.

	[image: image11.jpg]Percent Helium

LOW BATTERY

9N Teledyne
Helium/Oxygen Analyzer





	[image: image12.jpg]





16. Clean the Analyser with isopropyl alcohol to remove and marks and lightly polish.

17. Attach a “Vandagraph” sticker within the distributors’ box in the manual.

18. Generate stock sheets for the instrument and O2 sensor and ‘pass’, initial & date having tested satisfactorily.

19. Fit the UK pin adapter to the mains power supply and test for approximately 6.0V DC.

20 Supply 2lpm Flow restrictor  Pt.No.  


VM3COP54.01 Helium Analyser Test Procedure 050204.doc   26/10/2006      Page 4 of 4

