Vandagrapgh 

Replacement of the Alpha-1 Oxygen sensor

Using

Teledyne R-22DHO
Serial No’s: 194 and above

1) Remove the rear cover by undoing the 4 screws.

2) Disconnect the battery

3) Carefully withdraw the sensor connector lead from the sensor

4) Unscrew the Sensor from the case. The sensor may be tight  due to the sealing O ring and a tool such as a pair of large pliers or grip may be required  to rotate the sensor.

5) Position the new sensor inside the case with the screw thread aligned with the hole in the case. Gently screw the sensor into the hole taking care not to cross thread. The sensor thread is plastic and fine so can be easily damaged.
When the thread is finger tight the sensor can be gripped carefully with a tool and tightened gently just enough to compress the O ring seal.
Over tightening can damage either the thread or the O ring 

6) Replace the sensor connector, which can only be fitted one way. Ensure the connector is fully mated.

7) Replace the battery

8) Carry out a full calibration in accordance with the manufacturers instructions

The following notes are supplied as a guide only.
Carry out the calibration in dry outdoor air
Turn the front calibration control clockwise to maximum      
Adjust the internal calibration control until the display reads 22.2%
The setting 22.2% assumes an ambient pressure of 1013 mb
0.1% should be added for every increase of 5 mb and 0.1% subtracted for very 5 mB below
Seal the internal calibration control with a dab of paint

9) Replace the rear panel taking care not to over tighten the screws
10) Calibrate the outer calibration control to 20.9%.

Alpha sensor replacement  
Page 1
February 25, 2003

