QC 21 Non Conformance Report
	Date
	19/08/17

	Issue id
	101047

	BSI Ref (if applicable) 
	1511824-201709-M7

	Responsibility
	Derek Lamb

	Non-Conformance
	Area/process Pre-Clinical Data 
Clause ER 7.1 
Scope CE 01389 
Category 
Major 
Statement of 
nonconformance: 
The solutions adopted to fulfill the essential requirement 7.1 are incomplete. 
Clause 
requirements 
ER 7.1 – “The devices must be designed and manufactured in such a way as 
to guarantee the characteristics and performances referred to in Section I on 
the ‘General requirements’. Particular attention must be paid to: — the choice 
of materials used, particularly as regards toxicity…” 
Objective 
evidence 
Biocompatibility documentation does not fully meet requirements. Documents 
ID2474 and, ID7884 (No revisions provided) fail to: 
- Provide a Biological Evaluation Plan or overall Biological Evaluation Report 
made by a qualified person 
- Consider harmonized standards 
- Consider duration or nature of contact 
- Provided rationale for use of post market data. 
- Provide evidence of any consideration of biological impact. 
- Consider risk related to biological compatibility of the device.. 

	Investigation By:
	Derek Lamb

	Investigation Issue id
(if applicable)
	101047


	Corrective Action By:
	

	Corrective Action Issue ID (if applicable):
	101048


	Time Scale for Corrective Action
	

	Corrective Action:
	

	Follow-up future issue id
(Effectiveness verification)
	101049


	Effectiveness verification
	

	Closed By:
	

	Closed on
	


