QC 21 Non Conformance Report
	Date
	19/08/17

	Issue id
	101053

	BSI Ref (if applicable) 
	1511824-201709-M5 

	Responsibility
	Derek Lamb

	Non-Conformance
	Area/process Pre-Clinical Data 
Clause [Annex I ER 2] 
Scope CE 01389 
Category 
Major 
Assessment Report. 
Page 16 of 23 
Statement of 
nonconformance: 
The solution adopted to fulfil the Essential Requirement 2 is not complete. 
Clause 
requirements 
[Annex I ER 2] - The solutions adopted by the manufacturer for the design 
and construction of the devices must conform to safety principles, taking 
account of the generally acknowledged state of the art. 
Objective 
evidence 
There is no evidence of compliance to claimed standards, including EN 60601-2-10, and EN 60601-1. All clauses of the standard per report ID 11655 
(version 1332256919) have been listed explicitly as N/A including clearly 
applicable standards such as but not limited to section 7.0 and section 9.0. 
There is no evidence of compliance in part or in full to the harmonized 
standard EN 60601-1 or a rationale for partial compliance. 

	Investigation By:
	Derek Lamb

	Investigation Issue id
(if applicable)
	101053


	Corrective Action By:
	

	Corrective Action Issue ID (if applicable):
	101054


	Time Scale for Corrective Action
	

	Corrective Action:
	

	Follow-up future issue id
(Effectiveness verification)
	101055


	Effectiveness verification
	

	Closed By:
	

	Closed on
	


