QC 21 Non Conformance Report
	Date
	31/08/17

	Issue id
	101084

	BSI Ref (if applicable) 
	1515871-201705-M3

	Responsibility
	Derek Lamb

	Non-Conformance
	Area/process Risk Management File Clause ER2 Scope CE 540537 Category Major Statement of nonconformance: The solutions adopted to fulfill the essential requirements are incomplete. Clause requirements The solutions adopted by the manufacturer for the design and construction of the devices must conform to safety principles, taking account of the generally acknowledged state of the art Objective evidence The system presented systemically fails to follow ISO 147971 requirements, including but not limited to: The identified risk management document fails to (not limited to): - assignment of responsibilities and authorities per ISO 14971 section 3.4b - review of risk management activities per ISO 14971 section 3.4c - includes verification activities per ISO 14971 section 3.4e - includes activities related to production and post-production information collection and review per ISO 14971 sections 3.4f and 9 

While the risk management is recently updated there is no reference to post market surveillance data or evidence of maintenance of the risk management file in light of field data for the devices in question. 

Risks are not calculated both before and after risk controls (per 14971 section 6.4). 

Key relevant risks are identified as not applicable, with insufficient or incorrect rationales, including exclusion of monitoring functions as relevant risks.  
 

	Investigation By:
	Derek Lamb

	Investigation Issue id
(if applicable)
	101084
Risk management starts with a BI-Monthly issue to review ALL technical files,
Task50 See doc ID21516
VM3COP27.11 Performing a Technical File PMS and  risk assessment _ID17824
Page 6 of the BSI Report 8536454
"Findings 
Technical Documentation Assessed: 
MDD The assessment was based upon the Quality Management System, reference Quality Manual, 2017:20968 and a review of the Technical Documentation for the Pulse Oximetry (Pulse oximeter probes) Envitec Oxygen Sensors also called Pulse Oximetry Sensors 4000. "
All the Non Conformances reference CE 540537 
The Files sent for review Were relating to CE 97289
Page 6 Description is a Mix-up of both CE 540537  and CE 97289,
On the assumption the report is on the files provided, and are targeted to CE 97289, Envitec Oxygen Sensors.
exclusion of ‘monitoring functions’ as relevant risk is correct. 

	Corrective Action By:
	

	Corrective Action Issue ID (if applicable):
	101085


	Time Scale for Corrective Action
	

	Corrective Action:
	assignment of responsibilities docid


	Follow-up future issue id
(Effectiveness verification)
	101086


	Effectiveness verification
	

	Closed By:
	

	Closed on
	


