Waterproofing

The Microstim does not need waterproofing. It has been used for 30 years and we have no complaints, or records of damage by fluids. This is not surprising as the instrument is always used in a dry environment. In OR it will always be used outside of the sterile areain a dry clean area.

The Microstim should;d have an IP00 label affixed to it. To designate it is not dust or waterproof.

If third party testing is envisaged then mandatory waterproofing may be required
Useless though this may be the main ingress of fluid will be through the gaps shown as white in the photo
The rest of the enclosure appears waterproof. It should be possible to mold a cover that sits over the switches and can be glued to the switch and enclose. This drawing is not very good but shows the principle

The cover could be held in place with a rectangular washer and glued to the enclosure
The main reason for keeping the switches instead of a membrane is the Microstim is off when not producing pulses. This increases the battery life and prevents the instrument being left on after use.

Adding a display to the Microstim

The Microstim is a constant current device so it automatically compensates for any variations in patient impedance.
The method to determine the correct level of stimulation is for the instrument to be set at 1/3 rd position on the gain control. This has been established by multiple trial and error to locate a position which will not cause an undue shock to the patient.
The gain is then increased until the use sees the fingers twitch. Bio-feedbach controls the idea stimulation level. Adding a digital readout is purely cosmetic and will not improve the existing method and adds nothing but cost to the instrument. However if this is what the customer requires it can be done in several ways.

Using a generic circuit diagram NB Two capacitors have been moved on the final circuit.
TR1 & TR2 is a constant current generator, the level controlled by the voltage generated via RV1 and set by the base resistor . Pin 6 on the mode controller produces the stimulator pulses at 5v. It would therefore be possible to use the output from the VR1 and feed it into a Sample and hold circuit feeding a digital panel meter. This circuit would have to be activated when SW1 or SW2 are activated and would only read whilst the stimulator was in its active circuit.
Another solution is to disregard the actual stimulator cycle.
A separate circuit consisting of the equivalent to R5 RV1 and R11 could be fed whilst SW1 & SW2 are on . The DPM could be connected to RW2 would feed the DPM. RV1 & RV2 w2ould need to be mechanical linked and calibrated to match RV1

Assuming patient impedance is 1Kohms and the maximum output is 100v Then the current is 100MA

To set up each individual Microstim
Set output of 100v into 1Kohms and measure output voltage.
Set RV3 so DPM reads 100 MA. A small divider circuit may be required in front of the DPM

Settings of VR1 will be reflected in VR2 and the output current can be displayed. As this circuit does not require a patient to work the instrument can within the clinical parameters be pre prepared.
This solution will require a high grade high stability twin ganged potentiometer.

Conclusion
These solutions are additions to the original Microstim DB and would not infringe or invalidate the original CE mark.
As the circuit only adds an additional display and we Quote in the manual “ only an indication best method is bio-feedback “ the original specification and instructions are valid. Clinical trials should not be required

J.S.Lamb 18 Aug 2017