	     [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire BD20 7DT - United Kingdom
Tel: +44 1535 634542                 Fax: +44 1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



[image: image1.png]
Specification Datasheet – 0110182
Viamed

R-22AVG Oxygen Sensor
	    Specifications


	Part number
	0110182

	Model
	R-22AVG

	Output in ambient air
	9 mV – 13 mV 

	Output connector/connection
	3 pin Molex (22-11-1031)

	Input connector/connection
	M16 x 1 (5/8-24 UNEF)

	Input O ring
	Yes

	Measurement range
	0 – 100% oxygen (at atmospheric pressure)

	Repeatability
	< 1% volume oxygen at constant temperature and pressure

	Influence of pressure
	Output proportional to change in oxygen partial pressure

	Pressure range
	0.6 to 2.0 bar

	Linearity error
	< 3% relative

	Accuracy
	0 % to 3.3 % oxygen:     +/- 0.1 % absolute

3.4 % to 100 % oxygen: +/- 3.0 % relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 40µV in 100% nitrogen

	Cross interference
	≤ 0.1% volume oxygen response to: 

15% CO2 balance N2, 10% CO balance N2, 3000 ppm NO balance N2, 

3000 ppm C3H8 balance N2, 500 ppm H2S balance N2, 

500 ppm SO2 balance N2, 1000 ppm benzene balance N2

	Response time (t90)
	< 5 s for 90% of final value.

< 40 s fall time from 20.95% to 0.1% oxygen, when 100% nitrogen applied.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% relative after fall from 1m

	Operating temperature range
	0 to +50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC thermistor

	Temperature compensation error
	+25 to +40 ºC: 3% relative

    0 to +50 ºC: 8% relative

	Nominal sensor life
	≥ 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Recommended load
	≥ 10 K(

	Weight (approximate)
	28 g

	Flow diverter
	Not included

	Packaging
	Gas barrier bag – transparent.

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.

Specifications subject to change.



V1.0    08-06-2018

Page 1 of 1
[image: image2.jpg]


