	     [image: image3.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire BD20 7DT - United Kingdom
Tel: 01535 634542                            Fax: 01535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



0110021 R-22MEDV 

Long Term Performance Testing
Samples of production oxygen sensors are routinely retained for on-going long-term performance testing against the published specifications. To accelerate the testing, sensors are maintained in an environment of 100% oxygen, at room temperature. Sensors are routinely reprocessed through production final stage testing to ascertain when a sensor is exhausted or has failed. The following graph illustrates the lifetime results in terms of oxygen percent hours for recent sample sensors that have completed the testing process.
[image: image1.wmf]1

4

0

1

2

0

1

0

0

8

0

6

0

4

0

2

0

0

6

0

5

0

4

0

3

0

2

0

1

0

0

t

i

m

e

 

(

s

)

V

o

l

t

a

g

e

 

(

m

V

)

a

i

r

o

x

y

g

e

n

n

i

t

r

o

g

e

n


Oxygen sensors undergo the following final stage testing to verify that each sensor meets the published specifications in terms of: output in ambient air, output in oxygen, zero offset voltage, linearity error and response time.

Each sensor is exposed to a flow of dry ambient air for 10 seconds, followed by a flow of dry oxygen for 60 seconds and then to flow of dry nitrogen for a further 60 seconds. The flow rate is 3 L/min. 

A typical measurement output is shown in the figure.

[image: image4.png]Lifetime in % volume oxygen houers

Lifetime: storage in 100% oxygen Viamed 0110021
1600000
1400000
1200000
1000000
800000
600000

400000
200000

0
Apr12 Oct12 May13 Now13 Jun1d Decld JublS Jani6 Augls
Production date




Sensor parameters are derived as follows:

· Output in ambient air. 

During the last few seconds of the test, the sensor response is averaged and taken as the sensor output (mV). Actual real time barometric pressure is monitored and used to calculate the pressure compensated sensor output (mV) in ambient air at the standard pressure of 1013 hPa.
· Output in oxygen. 
Again, during the last few seconds of the test, the sensor response is averaged and pressure compensated as described above, so as to provide the sensor output (mV) in 100% oxygen.
· Zero offset voltage.  

The sensor response between 59.8 and 60 seconds in dry nitrogen is averaged and taken as the zero offset voltage.
· Linearity error.

The linearity error is calculated according to the following equation:


[image: image2.wmf]100

100

air

ambient

in

output

2095

.

0

oxygen

%

100

in

response

sensor

error

linearity

-

×

×

=


· T90 response time.

From the output in ambient air and the sensor response in 100% oxygen the response change is calculated. The time required to reach 10% and 91% of the response change respectively is taken to calculate the T90 time according to the following equation:

[image: image3.png]_1416985580

_1277885517.unknown

