Calibration Certificate Oxygen Analysers


Manufacturer:
Vandagraph Ltd.




15,Station Rd.,




Crosshills,




Keighley,




West Yorkshire,




BD20 7DT.




United Kingdom

Device:

VN202 Oxygen Analyser

Serial Number:

Description:

A Hand held portable Oxygen Analyser used to as a secondary device to 


measure Oxygen levels in mixed gases.

Calibration in Gases

Calibration in     100%                                                    Reading:

Exposure to Air   20.9%                                                  Reading:

Ratio   Reading 100%/ Reading in Air                          Ratio:                ( 4.764- 4.92)

Electronic Calibration:

Equipment; Oxycal  S/N                                                DVM   S/N

Set 50mV  Adjust Calibration to 100%

Voltage                     Expected Reading                      Reading  +/-1%      

50                               100

40                                 80

30                                 60

20                                 40

10                                 20

5                                   10








      Signature;








      Full Name;








      Position;

Certificate Number;








      Date ;

Calibration of Oxygen analysers & Monitors.doc                               26. November 2007


