Hospital:




Date:

Room:

Cabinet Body Serial number:
Check for overall signs of damage Comments

Check door catch – swivel catch on door

Check door catch retaining clasp on cabinet

Check integrity of welds on the door hinges

Check hinges for damage or deformation

Check the restraining bars for damage or deformation

Check the cabinet is securely fixed to wall

Check medirails and tighten any loose rails

Check grommets on the hose inlets are in place and intact

Check the bumper stops that prevent door banging during closing

Check caution sticker is in place on the platform

Check bed for damage i.e. cracks, chips

Check all the bed mounting screws are in place

Check mattress for damage i.e. tears, split seams
Tom Thumb – Serial Number:

Minimum reading less 7.85 mbar:

Adjustable valve 42.1 mbar +/-1:



Precision Valve 44.1 mbar +/-1:



Gauge @30cmh2o (28.4-30.4 mbar):
Check oxygen hose, if present, for wear, and replace if necessary (at least every 4 years)
Date of last hose replacement:
Low Suction Controller – Serial number:
Change filter assembly

Check vacuum hose for wear, and replace if necessary (at least every 4 years)
Date of last hose replacement:
Check rail clamp is securely attached to suction controller

Replace suction tubing

Replace tapered connector

Check receiving canister for damage, replaced?
Radiant Warmer Serial number:
Is Radiant Warmer hard-wired?

PAT test performed?
APGAR Timer – Serial number:
Change batteries

Check storage bins present

Tidy any trailing hoses, using tie-wraps if necessary

Check Operator’s Manuals are in place

Clean cabinet with isopropyl alcohol
Blender Serial number:
Date of last hose replacement:
@20.9:
@30:

@60:

@90:

@100:       (+/-3)

Check Alarm for no Oxygen

Check Alarm for no Air

Check O2/Air Balance Alarms before output change

Check Duck Bill Valves

Check for Leaks
