QC 21 Non Conformance Report
	Date
	Created:  31/10/17
Updated: 22/11/17

	Issue id
unique identifier 
	106959 

	BSI Ref (if applicable) 
unique identifier 
	1548900-201710-M1 

	Responsibility
Person Overall responsible 
	John Lamb

	Non-Conformance
statement of the problem 
	Design and development and risk management is not effective because evidence of a controlled process could not be determined.

1. The risk management process VM3COP27.11 includes a scale of 1-4 for severity and occurrence, but does not identify how to determine the boundary/limits between scales.

2. Evidence of consideration of Risk for the Tom Thumb design change in 2005 could not be provided at this assessment.

3. Evidence of updating the Tom Thumb design inputs for a design change in 2005 could not be provided at this assessment. (ref Tom Thumb specification document #2247 dated 6 June 1997)

4. Evidence of Tom Thumb design inputs containing the outputs of risk management could not be provided at this assessment (ref Tom Thumb specification document #2247 dated 6 June 1997 and Tom Thumb risk management file dated 29/9/2017)

5. Evidence of performing a design review on the Tom Thumb design change in 2005 could not be provided at this assessment.

6. Evidence of performing validation of the design change to Tom Thumb in 2005 could not be provided at this assessment.

7. Evidence of appropriate controls related to responsibility authority for design and development between VST and the subcontractor could not be provided at this assessment (ref contract #13859)

 

	Investigation By:
Person responsible 
	Derek Lamb


	Investigation Issue id
(if applicable)
Root Cause Analysis 
	106959 
1. Documents do not fully comply with the latest ISO13485; 2016. 

The documents had not been upgraded completely to ISO13485:2016. Process of upgrading and incorporation is on going. Any delay is primarily due to illness and three concurrent audits of Technical files.

2. There has been a transition period between paper files and the digital system. Intrastats Digital Archive was started 14 August 2006.

 When the technical files were originally scanned and brought into the system in August 2006. Not all the old records were added from the Goldmine system and paper files.
Document ID # 2994 found in the technical file (Y 14 Design change Rationale). After the BSI Audit so although available it was not easily locatable. 

3. There has been a transition period between paper files and the digital system. Intrastats Digital Archive was started 14 August 2006.

 When the technical files were originally scanned and brought into the system in August 2006. Not all the old records were added from the Goldmine system and paper files.
Document ID # 2994 found in the technical file (Y 14 Design change Rationale). After the BSI Audit so although available it was not easily locatable. As per 2.

4. The design files of Tom Thumb: circa 1993/4 were collated retrospectively into our technical files. Original designers are now retired or deceased and original files no longer available.

Please refer to the MDD and NB-MED/2.5.1/Rec5 

Page 2. 

5. There has been a transition period between paper files and the digital system. Intrastats Digital Archive was started 14 August 2006.

 When the technical files were originally scanned and brought into the system in August 2006. Not all the old records were added from the Goldmine system and paper files.
Document ID # 2994 found in the technical file (Y 14 Design change Rationale). After the BSI Audit so although available it was not easily locatable. As per 2.

6. There has been a transition period between paper files and the digital system. Intrastats Digital Archive was started 14 August 2006.

When the technical files were originally scanned and brought into the system in August 2006. Not all the old records were added from the Goldmine system and paper files.

Document ID # 2994 found in the technical file (Y 14 Design change Rationale). After the BSI Audit so although available it was not easily locatable. As per 2.

7. Discussions regarding the Scope of VST ISO 9001:2015 have been ongoing, due to the complex nature of VST, in general and where it sits in the design, development and manufacturing processes.

The primary root cause was VST was initially set up by cloning the Viamed system and as a result we cloned the subcontractor contract.



	Corrective Action By:
Person responsible 
	Derek Lamb / John Lamb


	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	

	Time Scale for Corrective Action
Time for completion of all identified actions 

	Immediate Action: 28th November 2017
Corrective Action: 31st December 2017

	Corrective Action:
	1.   
Immediate Action 
ID7393 ‘Risk Analysis Policy Definitions’ is being re-written in line with the current risk management process, with attention to EN 14971:2016.
ID7393 to be update and the Risk Management Policy and Definitions upgraded to reflect the requirements of EN14971,  MEDEV 2 12-1 rev 8 and ISO13485:2016. 
This will be applied into an Intrastats Risk analysis document.

Corrective Action
Risk boundary limits, per scale 1-5, need updating in each technical file individually and address probability scales, to include the approach to routine component failure. 

Document ID7393, located in Intrastats Document Index, needs linking to each technical file. This only uses levels 1-4, risk management process, has in line with newer expectations, been changed to a Levels of risk 1-5.

The definitions of Hazard severity and Level of risk will be amended to include 5 levels and defined.

Risk boundary limits need updating in each technical file individually.

Section 7 Design and Development will be been completely re-evaluated in line with the new ISO13485:2016.

2. 
Immediate Action
Document ID 23703 will be added to the system which includes communications with the MHRA, regarding the customers device request in 2005. 
With reference to design change in 2005, Steve Hardaker is to export the required information into Intrastats Technical Files. 
Further review of the archives will be done if required.
Corrective Action
As a result of the major non conformance, the current Technical Files section Y (Design Inputs) and Section Z (Design output) will be summarized and simplified into combined documents, primarily for the use by 3rd parties.
File Y15 – Validation, is to be extended. It will combine our present Design and Development Compliance with the present Project Validation, so that references to archived documentation can be sign-posted.
File Y14 – Design Changes, are to incorporate the -Design Changes Review. Which will include EN ISO 13485 section 7.3.9 (a-d).
A third party organisation, highly qualified and highly recommended by our Trade organisation, will be consulted on our current interpretation and implementation of our EN13485, MDD and future MDR with particular reference to historic products.
3. As per 2.
Immediate Action
Document ID 23703 will be added to the system which includes communications with the MHRA, regarding the customers device request in 2005. 
With reference to design change in 2005, Steve Hardaker is to export the required information into Intrastats Technical Files. 
Further review of the archives will be done if required.
Corrective Action
As a result of the major non conformance, the current Technical Files section Y (Design Inputs) and Section Z (Design output) will be summarized and simplified into combined documents, primarily for the use by 3rd parties.
With reference to design change in 2005, Steve Hardaker is to export any further required information into Intrastats Technical Files. 

File Y15 – Validation, is to be extended. It will combine our present Design and Development Compliance with the present Project Validation so that references to archived documentation can be sign-posted.

File Y14 – Design Changes, are to incorporate the -Design Changes Review. Which will include EN ISO 13485 section 7.3.9 (a-d).

A third party organisation, highly qualified and highly recommended by our Trade organisation, will be consulted on our current interpretation and implementation of our EN13485, MDD and future MDR with particular reference to historic products.

4.
Immediate Action
Design files of Tom Thumb: circa 1993/4 were collated retrospectively into our technical files. Original designers are now retired or deceased and original files no longer available. Please refer to the MDD and NB-ED/2.5.1/Rec5 Page 2.

Corrective Action
It is impossible to produce design files retrospectively for historic products. 

All we can do is ensure relevant historic documents are brought in to the new system when required. 

5.  Immediate Action 
Document ID 23703 to be added to the system.

Corrective Action
It is impossible to produce, any further design files, retrospectively for historic products. 

All we can do is ensure relevant historic documents are brought in to the new system when required. 

6.  Immediate Action
Document ID 23703 to be added to the system.
Corrective Action
The complete Design and development sections need investigation and upgrading where required. 

It is impossible to produce, any further design files, retrospectively for historic products. 

All we can do is ensure relevant historic documents are brought in to the new system when required. 

7. Immediate Action
A final decision to be made with regard VST Scope. 

Corrective Action
We now intend to limit the scope to the supply of oxygen sensors.

After All the above Corrective Actions:
The full updating of section 7 Design and development is as follows:

Section 7 Design and Development will be completely re-evaluated in line with the new ISO13485:2016.

VM3COP Risk Management Policy and definitions to be upgrade to reflect the requirements of EN14971. MEDEV 2 12-1 rev 8 will be applied into an Intrastats Risk analysis document.

The definitions Hazard severity and Level of risk to be amended to 5 levels and defined.

7.3.2a. QC22 Design Job Specification Brief to be expanded to cover all the requirements of the MDD Essential requirements including pre-start risk assessments. 

7.3.2b QC28 to be expanded to review reason for change, proposed change, action to be taken, personnel involved and associated risks.

7.3.2c Design Inputs versus Design outputs will be expanded to compare them against the original Design Job Specification Brief, with comments and samples tested, documented.

7.3.2d. Doc ID documents 7742 documents the people responsible and can be correlated against training history.

7.3.2e. As 7.3.2c. Design Inputs versus Design outputs will be expanded to compare them against the original Design Job Specification Brief, with comments and samples tested, documented.

7.3.2f The resources needed is stipulated in VOP02.

7.3.3a QC22.

7.3.3b QC22.

7.3.3c QC22.

7.3.3d As appropriate.

7.3.3.e As appropriate.

7.3.4a Design Inputs versus Design outputs.

7.3.4b U1 Purchase.

7.3.4c.  Design Inputs versus Design outputs, also M1 Packing validation.

7.3.4d.  Design Inputs versus Design outputs.

7.3.5a. QC28B.

7.3.5.b QC28B.

7.3.5.c QC28B.

7.3.6 QC30 Project Verification and Validation will be expanded to include Comments, Verification Doc ID, the number tested and the number passed. The sample size will be decided by the volume processed. Very small quantities or mass production, small quantities are individually subjected to QA.

7.3.8 Documents available for transfer from design to production.

M4 Drawings

J9 Sub assemblies and Circuit diagrams

J11 Photographs of sub assemblies

J6 Test methods

M5 Circuit Diagrams

N4 QA procedures

N5 Special processes

N6 Sub contractors

R1 Specifications of materials

U1Purchase specifications

VOPs as required

7.39 QC28B



	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


�	obviously vary on a case-by-case basis, depending on the type of product, the risk associated with its manufacture, installation, use and servicing, and the period that it has been on the market. For example, it is unlikely that well established products, regardless of their classification, will have much formal design validation – but the manufacturer will have considerable market experience of use, together with details on how he has responded to any problems that have emerged, and this available data should be used by way of validation.





8

