QC 21 Non Conformance Report

	Date
	Created:  31/10/17

Updated: 08/11/17

	Issue id
unique identifier 
	106962 

	BSI Ref (if applicable) 
unique identifier 
	1548900-201710-N1

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	The statement of the problem:

Actions in response to nonconforming product detected after delivery is not fully effective because evidence of review for vigilance is not clearly documented for nonconformities received which are not identified to be complaints.

Not all negative feedback nonconformities raised on Intrastat are determined to be complaints. A determination related to vigilance was stated to be made for these nonconformities, but not clearly documented in the record.

ref microstim returns.pdf document #19655



	Investigation By:
Person responsible 
	Derek Lamb

	Investigation Issue id

(if applicable)
Root Cause Analysis 
	106962
Microstim returns.pdf document #19655

is the resulting document from the yearly post market Surveillance reports on on products.

A similar non conformance has been raised from the technical files review.



	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	

	Time Scale for Corrective Action
Time for completion of all identified actions 
	03/11/2017

	Corrective Action:
	Due to an existing technical file review corrective action has already been started.

Immediate Action:
   All non conformance/Customer Complaints Issues now state in the data entry / review screens that “ *** NON Conformance Header - it must contain reasoning to explain why this is or is not a vigilance ISSUE *** ”
 The post market output report has been modified to include a more clear description of the header issues are related to – making it more obvious in the report if they are non conformance / product feed back headers or normal processing headers.
 Also added to the post market output  report :

Full output, including notes of all issues, after the summary pages. These will include the ‘determinations related to vigilance’

VOP 10 To be updated to include 

“ALL NON CONFORMANCE ISSUES must contain reasoning to explain why the non conformance is or is not a vigilance issue.”

VOP 19 To be Updated to included

“All Complaints Issues must contain reasoning to explain why the complaint is or is not a vigilance issue.”


	Follow-up future issue id

(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


