Specification list O2 sensor projects

Revision:	2018-05-02

Project:	21% O2 measurement:

- O2 cell in combination with electronics board (including interface) as an OEM product

- Dallas chip to limit the sensor life time; Dallas chip on a separate board attached to the sensor

- Sensor: 50-60 €, 30-40 € for the board

- Oxygen leak alarm system; alarm above 21% O2

- Interface type: RS232, analog as an option

- 5 pins from the Dallas chip board to the main PCB

| Specifications |  | Remarks |
| --- | --- | --- |
| Galvanic Solid State Sensor |  |  |
| ## Part number |  |  |
| ## Model number |  |  |
| Output in ambient air | 20 µA/ 21% O2 |  |
| Power requirement |  |  |
| ### Output connector | Existing sensor PCB, with 3 pin connection. |  |
| Input | Open input |  |
| Operating range | 0 – 30% oxygen (lower limit?) |  |
| Resolution | < 1% O2 |  |
| Pressure range | Ambient pressure |  |
| Repeatability | In the range of the resolution | Does this depend upon constant temperature and pressure? |
| Linearity error | ? e.g. 0.1 to 1.0 bar < ?% relative |  |
| Accuracy at constant temperature, pressure and humidity | Accuracy – Within + / - ?% of reading | Hysteresis accuracy? |
| Calibration interval | 3 months? | Ambient air calibration |
| Warm up time | < 1 minutes after sensor installed | Calculate the size of super cap to keep the cell supplied when mains is disconnected |
| Zero offset voltage | ? |  |
| Cross interference | ? | Anesthetic gases? |
| Response time | < 2 s for 90% of final value |  |
| Operating humidity | 5 – 99% R.H. | Condensation on the glass capillary must be prevented |
| Influence of humidity | ? % relative per % R.H. at 25 ºC |  |
| Influence of mechanical shock | < ? % after fall from 1m |  |
| Operating temperature range | e.g. 0 – 50 ºC |  |
| Storage temperature | e.g. -20 to +50 ºC | 60°C possible? |
| Recommended storage temperature | e.g. +5 to +15 ºC |  |
| Temperature compensation | Is this required? | Temperature compensation on the PCB possible/needed? |
| Temperature compensation error (steady state) | 0 °C to 50 °C ? % relative error |  |
| Nominal sensor life | 12 months |  |
| Typical sensor life | 12 - 24 months from date of invoice (including shelf life). Limited to 24 months by on-board integrated circuit. Ensures optimal linearity performance and integrity of the capillaries. |  |
| Long term output drift | < ? % volume oxygen per month, typically < ? % over sensor lifetime. |  |
| Load required | no | Direct load to cell not required, part of driving circuit |
| Applied standards | None | Alarm will be triggered from host system |

Project:	100% O2 measurement:

- Replacement for existing galvanic cells

- Analog output in the range of 9-13 mV @ ambient air

- Dallas chip to limit the sensor life time; Dallas chip on a separate board attached to the sensor

- OEM specific protection code possible

- 5 V supply voltage for the board

| Specifications |  | Remarks |
| --- | --- | --- |
| Galvanic Solid State Sensor |  |  |
| ## Part number |  |  |
| ## Model number |  |  |
| Output in ambient air | 3-5 µA @ 21% O2 |  |
| Power requirement |  |  |
| ### Output connector | Existing sensor PCB, with 3 pin connection. |  |
| Input | Open input |  |
| Operating range | 10 – 100% oxygen (lower limit?) |  |
| Resolution |  |  |
| Pressure range | e.g. 600 to 2500 mbar |  |
| Repeatability | <? % volume oxygen. | Does this depend upon constant temperature and pressure? |
| Linearity error | ? e.g. 0.1 to 1.0 bar < ?% relative |  |
| Accuracy at constant temperature, pressure and humidity | Accuracy – Within + / - ?% of reading | Hysteresis accuracy? See standard. |
| Calibration interval | 3 months? | Ambient air calibration |
| Warm up time | < ? seconds/minutes after sensor installed | Calculate the size of super cap to keep the cell supplied when mains is disconnected |
| Zero offset voltage | < ?% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds |  |
| Cross interference | < ?% volume oxygen response to 5% CO2 balance N2 | Anesthetic gases? See standard. |
| Response time | < 2 s for 90% of final value. |  |
| Operating humidity | 5 – 99% R.H. | Condensation on the glass capillary must be prevented |
| Influence of humidity | ? % relative per % R.H. at 25 ºC. |  |
| Influence of mechanical shock | < ? % after fall from 1m |  |
| Operating temperature range | e.g. 0 – 50 ºC |  |
| Storage temperature | e.g. -20 to +50 ºC | 60°C possible? |
| Recommended storage temperature | e.g. +5 to +15 ºC |  |
| Temperature compensation | Is this required? | Temperature compensation on the PCB possible/needed? |
| Temperature compensation error (steady state) | 0 °C to 50 °C, ? % relative error |  |
| Nominal sensor life | 24 months |  |
| Typical sensor life | 24 -36 months from date of invoice (including shelf life). Limited to 24 months by on-board integrated circuit. Ensures optimal linearity performance and integrity of the capillaries. | Shelf life needs to be discussed regarding Dallas chip (O2 operating hours?) |
| Long term output drift | < ? % volume oxygen per month, typically < ? % over sensor lifetime. |  |
| Load required | no | Direct load to cell not required, part of driving circuit |
| Applied standards | ISO 80601-2-55 |  |