Specification list O2 sensor projects

Revision:	2018-05-23

Project:	21% O2 measurement (alarm unit):

- O2 cell in combination with electronics board (including interface) as an OEM product

- Dallas chip to limit the sensor life time; Dallas chip on a separate board attached to the sensor

- Sensor: 50-60 €, 30-40 € for the board

- Oxygen leak alarm system; alarm above 21% O2

- Interface type: RS232, analog as an option

- 5 pins from the Dallas chip board to the main PCB

| Specifications |  | Remarks |
| --- | --- | --- |
| Galvanic Solid State Sensor |  |  |
| ## Part number | 01-ES1-O2-25-01 (EC Sense) | Will be defined later by Sensatronic |
| ## Model number |  | Will be defined later by Sensatronic Working title: <polymer micro oxygen sensor> |
| Output in ambient air | 20 µA/ 21% O2 | 40-60µA @ 21 Vol.-% O2 |
| Power requirement |  |  |
| ### Output connector | Existing sensor PCB, with 3 pin connection. |  |
| Gas Inlet | Open input, to ambient air/test gas |  |
| Operating range | Operation: 10 – 30% oxygen Calibration: 0 – 30% oxygen | 0 Vol-% O2 are for calibration are ok. Please consider the specs for the suitable humidity range. We expect no issues for prolonged exposure to 0 Vol.-% O2 if humidity is within specs. |
| Resolution | < 1% O2 |  |
| Pressure range | Ambient pressure (600 to 1200 hPa) | Ambient pressure effect is lower than with galvanic cells, but can be compensated by barometric pressure sensor |
| Repeatability | In the range of the resolution or less | Does this depend upon constant temperature and pressure? In general, yes! |
| Linearity error | < 2% of full scale | Response is linear, Error is currently missing in the datasheet. |
| Accuracy at constant temperature, pressure and humidity | Accuracy – Within + / - ?% of reading | Hysteresis accuracy? |
| Calibration interval | 3 months? | Ambient air calibration, triggered by vent with sending a status bit (real time clock needed?) |
| Warm up time | < 1 minutes after sensor installed | Calculate the size of super cap to keep the cell supplied when mains is disconnected |
| Zero offset voltage | <0.3 µA @ 0% oxygen |  |
| Cross interference | ? | Anesthetic gases? |
| Response time | < 2 s for 90% of final value | Should be comparable to galvanic cells |
| Operating humidity | 5 – 95% R.H. | Condensation on the glass capillary must be prevented, For long term exposure the limits must be change at least to 10 – 95% rH. Short term exposure ( e.g. <1h) is typically no problem. If there are applications which touch this operating conditions, we should discuss that cases. |
| Influence of humidity | ? % relative per % R.H. at 25 ºC |  |
| Influence of mechanical shock | < ? % after fall from 1m |  |
| Operating temperature range | e.g. 0 – 50 ºC |  |
| Storage temperature | e.g. -20 to +50 ºC | 60°C possible? Need to be checked if critical. Presumably possible in a sealed package for limited time (“Trunk test”) |
| Recommended storage temperature | e.g. +5 to +25 ºC | We store the sensors at room temperature. But it won’t hurt them if stored under these conditions if customers are used to. |
| Temperature compensation | Is this required? | Temperature compensation on the PCB possible/needed? Temperature of the cell is the critical point. |
| Temperature compensation error (steady state) | 0 °C to 50 °C ? % relative error | Need to be determined in the device. |
| Nominal sensor life | 12 months |  |
| Typical sensor life | 12 - 24 months from date of invoice (including shelf life). Limited to 24 months by on-board integrated circuit. Ensures optimal linearity performance and integrity of the capillaries. | EC-Sense: 24 Month warranty, storage life 6 month |
| Long term output drift | < 1 % of reading per month, typically < ? % over sensor lifetime. | Specs: <1% Signal / month ~ 0.2 Vol.-% / month at 21 %Vol. In practice lower values have been observed. |
| Load required | no | Direct load to cell not required, part of driving circuit |
| Applied standards | None | Alarm will be triggered by the host system |

Project:	100% O2 measurement (medical unit):

- Replacement for existing galvanic cells

- Analog output in the range of 9-13 mV @ ambient air

- Dallas chip to limit the sensor life time; Dallas chip on a separate board attached to the sensor

- OEM specific protection code possible

- 5 V supply voltage for the board

| Specifications |  | Remarks |
| --- | --- | --- |
| Galvanic Solid State Sensor |  |  |
| ## Part number |  |  |
| ## Model number |  |  |
| Output in ambient air | 3-5 µA @ 21% O2 |  |
| Power requirement |  |  |
| ### Output connector | Existing sensor PCB, with 3 pin connection. |  |
| Input | Open input |  |
| Operating range | 10 – 100% oxygen (lower limit?) |  |
| Resolution | < 1% O2 |  |
| Pressure range | e.g. 600 to 2500 mbar | Specs 800-1200hPa, 600hPa-1100hPa should be fine. Pressure compensation may be necessary. |
| Repeatability | <? % volume oxygen. | Does this depend upon constant temperature and pressure? |
| Linearity error | ? e.g. 0.1 to 1.0 bar < ?% relative | Response is linear, Error is currently missing in the datasheet. To clarify: What does this mean? 0.1 = 10% of the current oxygen concentration? One needs to distinguish two fundamental different use cases: 1) Work at “constant” absolute pressure and 2) Work with different absolute pressure. Even if the partial pressure is the same, the result may be different. That is what we like to test with the pressure chamber. |
| Accuracy at constant temperature, pressure and humidity | ± (volume of 2.5% O2 + 2.5% of gas concentration) (required by ISO 80601-2-55 Accuracy – Within + / - ?% of reading | Hysteresis accuracy? See standard. |
| Calibration interval | 3 months? | Ambient air calibration |
| Warm up time | < 1 minutes after sensor installed | Calculate the size of super cap to keep the cell supplied when mains is disconnected |
| Zero offset voltage | < ?% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds |  |
| Cross interference | < ?% volume oxygen response to 5% CO2 balance N2 | Anesthetic gases? See standard. |
| Response time | < 2 s for 90% of final value. |  |
| Operating humidity | 5 – 99% R.H. | Condensation on the glass capillary must be prevented, For long term exposure the limits must be change at least to 10 – 95% rH. Short term exposure ( e.g. <1h) is typically no problem. If there are applications which touch this operating conditions, we should discuss that cases. |
| Influence of humidity | ? % relative per % R.H. at 25 ºC. |  |
| Influence of mechanical shock | < ? % after fall from 1m |  |
| Operating temperature range | e.g. 0 – 50 ºC |  |
| Storage temperature | e.g. -20 to +50 ºC | 60°C possible? Need to be checked if critical. Presumably possible in a sealed package for limited time (“Trunk test”) |
| Recommended storage temperature | e.g. +5 to +15 ºC | We store the sensors at room temperature. But it won’t hurt them if stored under these conditions if customers are used to. |
| Temperature compensation | Is this required? | Temperature compensation on the PCB possible/needed? Depending on the accuracy requirements and environment |
| Temperature compensation error (steady state) | 0 °C to 50 °C, ? % relative error | Need to be determined in the device. |
| Nominal sensor life | 24 months |  |
| Typical sensor life | 24-36 months from date of invoice (including shelf life). Limited to 24 months by on-board integrated circuit. Ensures optimal linearity performance and integrity of the capillaries. | Shelf life needs to be discussed regarding Dallas chip (O2 operating hours?) EC-Sense: 24 Month warranty, storage life 6 month |
| Long term output drift | < 1 % volume oxygen per month, typically < ? % over sensor lifetime. ISO 80601-2-55: stable based on accuracy requirements within 6 h of operation | Specs: <1% Signal / month ~ 0.2 Vol.-% / month at 21 %Vol. In practice lower values have been observed. |
| Load required | no | Direct load to cell not required, part of driving circuit |
| Applied standards | ISO 80601-2-55 |  |