
Showing minutes for : 
Held on 19 Sep 2007 09:09:55
(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Oxygen sensor bag and bag label.


19 Sep 2007 Derek Lamb
Issue to: Changed from General To Steve Nixon
Re-labeling of sensors:
Req.

JSL: confirmed Italy have demanded the Use by date on sensors:

EAN Standard states we need a use by date: 13.3E states

Kunal confirms we will get an answer by next week.

Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Automotive sensors output 9-13mv 

19 Sep 2007 Derek Lamb
Issue to: Changed from General To Steve Nixon
SN machines are now rejecting sensors,
IT / envitec already produce these sensors

kunal: no cherry pickng of sensors:
R-17BOS - 7-13Mv output.

Kunal will find out about the situation of 9-13Mv
will come back to use next week

SN clarififed not to change the AH`s
only involves R17V R21V R23V
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
R22Med with output 9-13mv 
< 5 s > 500,000 OOM103 

19 Sep 2007 Derek Lamb
OOM103 - Envitec number
less 5 S more than 500000Hrs
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
R17Med with Output 9-13Mv OOM103-1
Completion Notes
See Other Issues related to subject

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Barcoding and EAN bar codes 

19 Sep 2007 Derek Lamb
Teledyne currently looking into Barcodeing serialnumbers:

Someone in Teledyne will contact us with regard to registering products with EAN Barcodes
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
WEEE Label 

19 Sep 2007 Derek Lamb
Kunal:
Medical products are Currently Exempt from WEEE Directives

JSL:
UK Government are saying the Exemption status is Temporary,

SN: Batterys are coming under a different set of regualations by 2009, are sensors coming under the battery heading?

Kunal: O2 sensors not used in the same way as Battery and will stay exempt.


JSL / SN: As far as we are aware the WEEE logo has to appear on the AX/MX and GB300 oxygen monitors.

WE need in writing from Teledyne that we are allowed to stick the WEEE logo onto the Device.
As per the DTI website WEEE does apply to medical devices

http://www.berr.gov.uk/innovation/sustainability/weee/page30269.html
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Revised MX300 Software 

19 Sep 2007 Derek Lamb
Kunal: all complete and done - awaiting final Q.A.

Will come back to us to find out what the hold up is:



19 Sep 2007 Derek Lamb
All new units to come with new software
Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
AX300 Software - Ref Alarm Button 

19 Sep 2007 Derek Lamb
Issue to: Changed from General To Steve Nixon
no customers have reaised this as an issue but Viamed located the potential problem.

If AX alarm silence button sticks.
unit only recognises a sticking button in the first 30second of turning the unit on,
after which a stuck button is not detected - allowing for permenant silenced mode.
- unit should produce an error code on screen if the silence button sticks
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
MOX-3 Required T7 Based
Completion Notes
See Sensor modification Issue:

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Accutronics Fabian 9-13.5mV R22 Body with 3 Pin Molex
Completion Notes
See Sensor modification Issue:

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Viasys Avea (sensor 62689), supplied with AI PSR-11-917-M, R22MED too wide. R22 Spec, R24 body. 

19 Sep 2007 Derek Lamb
See Sensor modification Issue:
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Blease Fabian Genius R125P04-001 Maxtec 

19 Sep 2007 Derek Lamb
See Sensor modification Issue:
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Require R22CC 9-13mV, <13s, >1,000,000 OOM102 

19 Sep 2007 Derek Lamb
See Sensor modification Issue:
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Envitec OOM104, OOM106, OOM202-1, OOM202-1, OOM202-2, OOM202-2S, OOM204 

19 Sep 2007 Derek Lamb
See Sensor modification Issue:
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Delay in supplying R22MEDs
Completion Notes
SN:
General delivery is better than it used to be

Returns system seems to be working.


Delay in supplying R22MEDs- was a problem in the past.

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Tight AX/MX cables affect readings
Completion Notes
Discussed

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Socket problems with AX/MX 

19 Sep 2007 Derek Lamb
Discussed
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
AX/MX translations 

19 Sep 2007 Derek Lamb
JSL:
We believe that its going to be mandatory that we produce manual in 28 Languages.
Completion Notes
SN: certain countrys now stating they want it in thier local language.
we may lose the ability to sell in various countrys.

Kunal: currently its mandatrory to translate into french german spanish english italian, thats what teledyne will do for the time being,

each other language to be looked at on a case by case basis.

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Plans for alternative or additions to AX/MX?
Completion Notes
kunal: Not currently.

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
AX/MX cable problems 

19 Sep 2007 Derek Lamb
Discussed
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Update of existing customers` MX300s 

19 Sep 2007 Derek Lamb
See 11579
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Update of existing customers` AX300s
19 Sep 2007 Derek Lamb
kunal: suggest board swap, we will ing to do the work,
is Teledyne happy to provide boards;

Answer expected within about 2 weeks
19 Sep 2007 Derek Lamb
aproximate number 50 units 

19 Sep 2007 Derek Lamb
Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Sensor serial number errors. 

19 Sep 2007 Derek Lamb
Issue to: Changed from General To Steve Nixon
Delivery note Different to actual product.

SN to bring to Teledynes notice the moment these errors occur
Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Consolidation of shipments 

19 Sep 2007 Derek Lamb
Issue to: Changed from General To Steve Nixon
Seperate shipments for vandagraph / viamed medical / anything hazardous.

R29 R29I R30 - if ordered direct from maxtec no notice of hazardous,
why when ordered from teledyne are they label hazardous?.

SN to locate possible solutions and report it if its possible to consolidate all the shipments.
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Incorrect front panels on AX/MX i.e. Green & % nitrogen
Completion Notes
how why?

Kunal: they were made for special customer, appears mistake was made:

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
AN300 developments?
Completion Notes
kunal: 3800 units sold worldwide:
south east asia : 3000 market size in first year


(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
GB300 how to proceed.
Completion Notes
Kunal:
we able to purchase and supply

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Process Analyser Systems – how to proceed?
Completion Notes
see early issue

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Spares for TED60...
Completion Notes
Issue #11626

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Spares for TED191...
Completion Notes
Issue #11626

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Spares for TED200...
Completion Notes
Issue #11626

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
TED 60 pot P-456
Completion Notes
Issue #11626

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
TED 200 keypads
Completion Notes
Issue #11626

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
HAZ incorrect statements on invoices
Completion Notes
discussed

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Shipping charges
Completion Notes
getting better,

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
China sales and IT
Completion Notes
discussed go with current stragergy (18$)

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
AN300 displays 100 without cable connected.
Completion Notes
Disscussed
not worth the update

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Zero offset errors 

19 Sep 2007 Derek Lamb
no answer yet from Tom 
we are loosing sales in italy due to this and giving the perception that teledyne are not top quality sensors
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
QA test results
Completion Notes
discussed

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
RS232 cables supplied either without or varying connectors. 

19 Sep 2007 Derek Lamb
discussed
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Envitec/IT pricing
Completion Notes
discussed

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
MAX-43 

19 Sep 2007 Derek Lamb
Issue to: Changed from General To Steve Nixon
new ohmeda for giraffee
bought in pairs
IT gambit makes,
Send sensor and specs to Kunal
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Vandagraph
Battery Compartments 

19 Sep 2007 Derek Lamb
Mixcheck:
hard to get batterys in.
depends on the mounting position of the battery holder
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Vandagraph
Springs Screws
Completion Notes
discussed

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Vandagraph
Damage Spring Screw stats
Completion Notes
discussed

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Vandagraph
Re-Label Sensors
Completion Notes
discussed under viamed same applies to vandagraph

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Vandagraph
Modify Sensors
Completion Notes
discussed under viamed same applies to vandagraph

(0.0) Teledyne Issues 

Issued To:General
Created: 18 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
18 Sep 2007 Derek Lamb
Vandagraph
Pricing
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Kunal.
54$ per sensor to france
20 pieces per year,

medical / diving side teledyne not selling products.
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Kunal to deal with PAS
http://www.o2-sensors.co.uk/reference.htm

not supposed to be selling diving or Medical sensors,

Kunal stated to PAS he should buy from Viamed.
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Kunal being asked about china/india markets for medical products

Jsl:
10-11$ to get the medical market in china,

SN
IT making sensors for 7-8$.

Jsl
Oxygen Monitors just starting to appear in china 

China Sensor spec are not 7-13Mv, ryan to return with sensor spec.
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Kunal:
India oxygen sensors big market - req.s 7-8 Distributors

Lux..? 1 person
Ahmedabad 1 Person
Babglor? 1 Person
cochin or madras? - 1 Person
pune? 1 person
bhopal? 1 person

each citys population 5mil.
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Kunal:
China oxygen sensors distributors
chenzin? 1 person
chengdo? 1 person
dushanzi? 1 person
beging? 1 person

JSL: prblem gettting small qtys into china
Looked at possiblinlty of importing the products into china and then to distribute, adds another layer of cost into the price

Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
SN: Issue for China is pricing.
2-3 years ago requested 16$
now have price 18$
envitec and IT have gone in with low pricing.
IT have introduced sensors that Teledyne currently dont have.



Kunal:
India medical products we should get premium price
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Kunal : Sensor WILL sell at a high price in India
Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Sensor Label Issue:
Can Teledyne put the viamed address on the sensor label.

Kunal can see the logic.

Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
Added by Derek Lamb
Send link to Kunal with link to specification page showing 9-13Mvfor sensors
Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Viamed requires a equivelant of R22CC
- OOM102 (envitecs)
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Modifing Sensors Teledyne will allow this (in Kunals view) providing we remove the teledyne name from the sensor,

may need terminology of what we put on the label approving if teledyne name to stay on the label,

Possible we may need to sign a `side` contract with Teledyne.

Kunal will work on the issue.
Concrete Plan deadline before 2008.

JSL: STATED Vandagraph will no modify any sensors destined for use in rebreathers.

Modifiying of medical sensors brings huge liability risks to Viamed.Needs properly reviewing.





Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
SN to produce for each uniquie sensors a detailed spec and a sales projection.
and send to Kunal,
done within the next few weeks, - data by end of October and resolve the whole issue by 1 jan 2008
Completion Notes

(0.0) Teledyne Issues 

Issued To:John Lamb
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
To Produce Specification of new VANDAGRAPH sensors and sales projections,
by end Oct ober 2007,
Plan to be in place by 1 Jan 2008
Completion Notes

(0.0) Teledyne Issues 

Issued To:Steve Nixon
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
Added by Derek Lamb
Are AX /MX ROSH compatible?
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Vandagraph
$275-300 Oxygen Analyser in Diving Market.
Water Proof essential.

Analox have world lead
Maxtec selling something similar

guesstimate a few thousand units per year - market size is growing.

Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Name of person in Trianin to be given to Steve Nixon so he can contact with regard to AX/MX cables
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Any spares of old units that can be scavanged are required
Completion Notes

(0.0) Teledyne Issues 

Issued To:General
Created: 19 Sep 2007
Completed 19 Sep 2007
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
kunal to supply the Output Specifications of automotive sensors,
of oems
Completion Notes

(0.0) Teledyne Issues 

Issued To:Derek Lamb
Created: 19 Sep 2007
Completed Still Outstanding
Stock Ref:
Opera A/c: 
Order # 
Notes
19 Sep 2007 Derek Lamb
Sensor test sheet format / layout to be provided to teledyne of what we would like
Completion Notes



�<!--StartFragment-->


�<!--EndFragment-->





