30th April 2018

User Manual queries and suggestions – Handheld Pulse Oximeter VM-2160

Manual used in conjunction with VM-2160 SMARTsat, installed with firmware:

version SCE V5.8 English

Page 2/28

Copyright 2010, should it be 2017?

HF: I thought the date does not change from first release of the document. Not correct?

Yes, but is this first one for VM-2160 SMARTsat or do we treat as a progression of VM-2160. I think leave as is.

Suggest change Distributor Contact information to Sales & Service contact information

HF: implemented Thanks

Page 5/28

Why have you stated insured/user ?

HF: was required in the German text to receive an official number from health insurance which allows easy billing of the device. But I removed it as it sounds strange in English…

Thanks

Page 6/28

Should induced current be induced electrical current ?

HF: implemented Thanks

Page 7/28

Suggest changing:

Multifunction buttons used for

- Scrolling through menu items and

- Increasing/decreasing parameters.

to

Multifunction buttons used for:

- Scrolling through menu items.

- Increasing/decreasing parameter values.

HF: implemented

Thanks

Page 8/28

- No mention of bar graph colour but covered later in the manual.

HF: added “The colour of the bar graph is an indicator for signal quality (refer to 2.3)” OK

- Suggest changing with strong to of good.

HF: implemented Thanks

- Suggest changing to Time-interval of trends

HF: implemented Thanks

Is the description for Display 3 to 5 valid?

Example shown of 15-minute trend display mode. Trend data is stored in parallel to on-going measurement.

Display 3 = 15-minute trend

Display 4 = 30-minute trend

Display 5 = 4-hour trend

HF: updated to be the same like for CapnoTrue  OK

Display 3 to 5: Trend

(15 min, 1 h and 6 h Trend, parallel to ongoing measurement)

Page 10/28

Suggest changing always overlay alarms to always override alarms

HF: implemented  Thanks

The high priority alarm appears to be with just a 1 second pause. Why have you stared two versions: 2 seconds pause and 3 second pause?

HF: indeed, it should be 1sec, thank you, great catch!!! Implemented correction OK

2 versions (1sec and 3sec) due to IEC_60601-1-8:

On the Low priority section suggest changing – Ambient light to – Excessive ambient light

HF: Not possible, too long to fit on display screen in monitor  OK, understood

Page 11/28

Suggest changing for a longer period of time to for a long period of time

HF: implemented Thanks

Page12/28

NOTE. In regard to the last warning, the single tone will not be heard if the loudspeaker has been switched off on the previous power down.

HF added description of start-up tone mute function in sections “2.1, 2.3, 3.3”.

Removed warning (not helpful, only confuses), instead added in section 3.4 “Ensure that the monitor is providing a reading and that the pulse tone can be switched on and off to verify proper operation.”  OK, thanks

Page 13/28

NOTE. 3.5 Under what circumstances is the device considered not in use? If sensor removed from site or disconnected, the unit will not switch off.

HF: added “, provided no alarm condition is present”. Does this clarify enough?

OK, but unless alarms are turned off, the unit will never switch off?

Page 14/28

NOTE. Visual alarm - Only the critical value flashes yellow, the violated alarm limit turns yellow but does not flash!

HF: implemented correction Thanks

Page 15/28

4.1.3 should be Submenu: Setup not Submenu: Device setup

HF: implemented correction Thanks

Page 17/28

Suggest changing episodes like over-night to episodes such as over-night

HF: implemented correction Thanks

Page 18/28

4.2.2 Power-Save mode.

Do you know that if you keep your finger pressed, the screen flashes between two display formats. This does not look correct?

HF: implemented correction Thanks

Page 19/28

References EN ISO 9919:2005 rather than 80601-2-61 ?

HF: implemented correction Thanks

Page 20/28

Change The monitor generates uses and radiates to The monitor generates, uses and radiates

HF: implemented correction Thanks

Page 23/28

Pulse rate is specified as 2 Arms, shouldn’t this be a value of bpm i.e. 2 bpm?

So change from

20-300/Min. +/- 2 Arms

To

20-300 bpm +/- 2 bpm

Do we need to state ‘no motion’?

HF: added (no motion). OK

According than 80601-2-61 need to also have Arms for pulse rate:

OK, didn’t know that. We should also update the manuals and leaflets for the OEM boards.

Page 24/25

IEC 60 601-1-8:2006 change to IEC 60601-1-8:2006

HF: implemented correction  Thanks

No mention of the following, which there have been previous reference to:

IEC 60601-1-2:2014 (4th edition)

IEC 60601-1-11:2010

HF: we don’t have Test reports for these  OK, but these are applied to the OEM boards?

Page 25/28

The pool consists of healthy, female and male with skin tones ranging from light to dark subjects that are 18 years of age or older.

Should be

The pool consists of healthy female and male subjects that are 18 years of age or older, with skin tones ranging from light to dark.

HF: implemented correction Thanks

At bottom of the page change (5 V DC )and to (5 V DC) and

HF: implemented correction  Thanks

Page 26/28

Could we add our order numbers for each version?

0012165 – Central European

0012166 - Scandinavian

0012167 - Special European Character

HF: implemented Thanks

Why is Dutch in the Scandinavian package?

Historically due to a tender OK

Page 27/28

Please change REF 0014751 to 0014754

HF: implemented Thanks

Should we state that the wrap sensor is soft silicone. What about patient weight range, cable length (1.2M),  cable material – silicone.

HF: implemented, however no weight indication as it depends on measurement site OK, understood

Should we add the sensor wraps for the wrap sensor, 0014890?

HF: implemented. Do you sell them single, or in a box?  In a box.

Can we add carrying case – soft pouch 0022178?

HF: implemented. Thanks

Should we add the cable lengths and cable material to the list of disposables?

Adult  0.45m cable, PVC

Paediatric  0.45m cable, PVC

Infant  0.90m cable, PVC

Neonatal  0.90m cable, PVC

HF: implemented. Thanks