Customer error description:

Bluepoint service error analysis:

| Name |  |
| --- | --- |
| Screw locking: yes No Software version: Power supply connector: male, black male, metal female, black Gas outlet: back top no CapnoTrue®ASP Analysis result: |  |

Actions for repair:

SW update to version: HW update: (regularly carried out from v5.9) (ASP gas outlet is regularly moved to top housing: Mod.4) 24h load test performed (regularly carried out if main PCBA is replaced) : yes no Actions for repair and replaced parts (Name, REF, SN):

Comment to customer:

| Test Result (detailed test see next pages) | Pass AMP/MG: Step 1-5+7, ASP: Step 1-7 |  | Fail Don’t use device to monitor patients! |
| --- | --- | --- | --- |
| Comments: |  |  |  |
| Date |  | Tester |  |

| Test step | Acceptance criteria |  | Result |  |  | comment |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1. Visual inspection |  |  |  |  |  |  |  |
| 1.1) Housing, Display, Keypad | No damages or defects |  | pass |  | fail |  |  |
| 1.2) Sensors and cable | No exposed optical components, cables or other damages to the sensor |  | pass |  | fail |  |  |
| 1.3) External power supply (cable+ connector) | - No damages or defects. - Output voltage: 6,0V - 6,2V. - Electrical test (e.g. SECUTEST VDE 0751) is passed. |  | pass |  | fail |  |  |
| 1.4) Label + transparent protective foil | Label and the transparent protective foil are fixed clean to the back of the device. |  | pass |  | fail |  |  |
| 1.5) Screw lock | Fixed to upper right-hand screw |  | pass |  | fail |  |  |
| 2. Power Supply Options (AA batteries, Li-Poly rechargeable battery, external power supply) Test procedure: switch on with each supply option (see test steps 2.1 - 2.5) >perform Measurements: SpO2 at finger; etCO2 by breathing through airway adapter > switch off |  |  |  |  |  |  |  |
| 2.1) Only 4xAA Batteries | Battery symbol is displayed. Proper operation (device is providing SpO2 & CO2 reading) |  | pass |  | fail |  |  |
| 2.2) Only Li-Poly | Battery symbol is displayed. Proper operation (device is providing SpO2 & CO2 reading) |  | pass |  | fail |  |  |
| 2.3) Only ext. power supply | Power supply symbol is displayed. Proper operation (device is providing SpO2 & CO2 reading) |  | pass |  | fail |  |  |
| 2.4) ext. power supply + Li-Poly | Charging symbol is displayed. Proper operation (device is providing SpO2 & CO2 reading). |  | pass |  | fail |  |  |
| 2.5) No sensors are connected to the device. Switch on with 4xAA Batteries with <4.1V. | Low battery alarm (visual and audible) |  | pass |  | fail |  |  |
| 3. SpO2-Function at simulator (Fluke Index 2, SpO2 = 82%, Pulse rate = 75bpm, R-curve=BPM) |  |  |  |  |  |  |  |
| 3.1) SpO2 at simulator set point: 82% | 82% SpO2 (+-2 digits) |  | pass |  | fail | reading SpO2: |  |
| 3.2) Pulse rate at simulator set point: 75bpm | 75 beats/min (+-1 digits) |  | pass |  | fail | reading PR: |  |
| 3.3) SpO2 low alarm (at default low limit setting: 85%) | Audible and visual |  | pass |  | fail |  |  |
| 4. CO2-Measurement (CO2 values at breath simulator: CapnoCheck Program 1, 10 breath per minute) Select Mode at CapnoTrue: Main Menu > Service > Maintenance > Check gas accuracy Calibration gas (5% ±0.1vol% CO2, 20.9% O2 balance N2) |  |  |  |  |  |  |  |
| 4.1) 5vol% CO2 calibration gas + airway adaptor | EtCO2:4.7 – 5.3 vol% FiCO2: 0.0 – 0.3vol% |  | pass |  | fail | value: EtCO2 / FiCO2 | Zeroing? yes no |
| 4.2) Breath per minute | 10 (+-1 bpm) |  | pass |  | fail | reading RR:: |  |
| 5. Alarms und Data Download (switch on device, SpO2 Sensor at finger, breath through airway adapter) |  |  |  |  |  |  |  |
| 5.1) Remove SpO2 Sensor from finger | Alarm sound and message „SpO2 Probe off !“ |  |  | pass | fail |  |  |
| 5.2) Breath 3x through airway adapter and wait 20s | Alarm sound and message „Apnoea !!“ |  |  | pass | fail |  |  |
| 5.3) Download data to PC with CapnoTrue PC-Software | Successful data download to PC |  |  | pass | fail | PC-SW version: |  |
| 5.4) Delete all data | All stored data in the device are deleted |  |  | pass | fail |  |  |
| 6. Leakage test – ISA module (Only for sidestream capnograph CapnoTrue®ASP) |  |  |  |  |  |  |  |
| 6.1) Perform leakage test (see Technical Support Manual) |  | CO2 drop <0.5vol% in 1 Min. | pass |  | fail |  |  |
| 7. Packaging and accessories |  |  |  |  |  |  |  |
| 7.1) Latest FW on the device? | See server |  | pass |  | fail | Ver. Model Language |  |
| 7.2) Time (accuracy RTC) | Same as test PC |  | pass |  | fail |  |  |
| 7.3) Remove Li-ion battery | Battery case empty |  | pass |  | fail |  |  |
| 7.4) SN on device back label | SN like label in battery case |  | pass |  | fail |  |  |
| 7.5) Protective sheet on label | Implemented |  | pass |  | fail |  |  |
| 7.6) Screw locking | Implemented |  | pass |  | fail |  |  |
| 7.7) Final check | Device and parts clean& no damages |  | pass |  | fail |  |  |

Attachments: