	 
[image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire  BD20 7DT - United Kingdom
Tel: +44 (0)1535 634542     Fax: +44 (0)1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk


V1000 Foetal Heart Simulator

Part Number: 1410000

Performance Evaluation
	Date
	 21/02/07
	

	Serial number
	 PR065A21
	

	
	Pass
	Fail
	

	On/Off control
	(
	
	

	Power - good
	Green LED
	(
	
	

	Power – low
	Green LED

‘flashing’
	(
	
	

	BPM control
	(
	
	

	
	BPM achieved
	
	
	
	

	BPM  120
	120
	Yellow LED
	(
	
	

	BPM  150
	150
	Yellow LED
	(
	
	

	BPM  180
	180
	Yellow LED
	(
	
	

	BPM  210
	210
	Yellow LED
	(
	
	

	BPM  Auto
	Auto
	Yellow LED
	(
	
	

	BPM  30
	N/A
	Yellow LED
	(
	
	

	BPM  60
	60
	Yellow LED
	(
	
	

	BPM  90
	90
	Yellow LED
	(
	
	

	Intensity control functionality
	(
	
	


Performance evaluation:  PASS 

Signed:    ______________________                                                                                         

Name:     Philip Crossley

_1122892293.bin

