Handheld Pulse Oximeter

OxyTrue® A / S

With ChipOx Technology

Service manual

OxyTrue® A/S (ChipOx) - Service Manual

REV: 3

Released: 11/2017

The information in this document is subject to change without notice.

Copyright © 2008 bluepoint MEDICAL GmbH & Co. KG. All Rights Reserved.

This document contains proprietary information that is protected by copyright. All rights are reserved. No part of this document may be photocopied, reproduced or translated to another language without prior written consent of bluepoint MEDICAL GmbH & Co. KG.

Contact information

bluepoint MEDICAL GmbH & Co. KG

An der Trave 15

23923 Selmsdorf

Germany

Phone:	 +49-(0)-38823-5488-00

Fax:     +49-(0)-38823-5488-29

E-mail: info@bluepoint-medical.com

www.bluepoint-medical.com

Table of Contents

1	Introduction	4

2	Service menu	4

2.1	Service PIN	4

2.2	Service menu structure	4

3	Maintenance	6

4	Service and repairs	6

4.1	General	6

4.1.1	Device disassembly	6

4.1.2	Device assembly	6

4.1.3	OxyTrue® A components and HW compatibility	7

4.1.4	OxyTrue® S components and HW compatibility	8

4.2	Replace Components	9

4.2.1	Main-PCB	9

4.2.2	USB-PCB	9

4.2.3	USB-PCB cable	10

4.2.4	Flat ribbon cable	11

4.2.5	SpO2 Module	11

4.2.6	SpO2 sensor socket	12

4.2.7	Keypad	13

4.2.8	Display window	13

4.2.9	Housing	14

5	Troubleshooting Guide	15

6	Spare parts	16

7	Document History	17

# Introduction

This service manual contains information for servicing the OxyTrue® A / S handheld pulse oximeter. Only authorized service personnel may service this product.

The Handheld Pulse Oximeter OxyTrue ® A / S User Manual (the operator’s manual) is an integral part of the service procedures. Before servicing the monitor, read the operator’s manual carefully for a thorough understanding of how to operate the unit.

Only use the tools and test equipment as specified in this manual. Only use original spare parts available from your local distributor.

NOTE: This manual is for OxyTrue®A/S with ChipOx SpO2 technology. A separate manual is available for the OxyTrue®A monitor with high performance SMARTsat® technology.

Warning:

Incorrect procedures may harm the patient, or damage the monitor.

# Service menu

The Service menu is accessible to authorized personnel only. The PIN Code specified below shall not be given to any third parties. Changes of settings are only allowed with profound knowledge about each setting.

Warning:

Wrong settings may lead to misinterpretation and may affect patient safety!

## Service PIN

The PIN Code for OxyTrue® A/S can be accessed by menu item “Service”.

By entering the Service PIN CODE you confirm that you have read and understood the service manual and take full responsibility for the consequences that might result from the changed alarm functionality.

PIN CODE :  1  9  8  5

## Service menu structure

Within the service menu point START UP SETTINGS selective default start up parameters can be adjusted or the default start up parameters can be reset to the factory defaults.

The service menu point SYSTEM INFORMATION is for factory internal use only; there are no adjustable parameters in this menu.

Warning:

Changing the default start up settings of the device can seriously affect the functionality of the device and the alarms.

Warning:

Only authorized personnel with the required technical and medical knowledge are allowed to change the default start up settings.

Start-up setting:

| Parameter | Possible settings | Factory default | Description |
| --- | --- | --- | --- |
| Volume | 1/ 2/ 3/ 4 | 2 | Set default loudness. 1 = off |
| Alarm SpO2 ↑ | 100 to Alarm SpO2 ↓+1/ Off | 100 | Set upper alarm limit SpO2 |
| Alarm SpO2 ↓ | 45 to Alarm SpO2 ↑-1/ Off | 85 | Set lower alarm limit SpO2 |
| Alarm Pulse ↑ | 300 to Alarm Pulse ↓+1/ Off | 140 | Set upper alarm limit pulse rate |
| Alarm Pulse ↓ | 20 to Alarm Pulse ↑-1/ Off | 45 | Set lower alarm limit pulse rate |
| High Alarms | On/Off | On | Turn alarms of high priority on/ off |
| Medium Alarms | On/Off | On | Turn alarms of medium priority on/ off |
| Low Alarms | On/Off | On | Turn alarms of low priority on/ off |
| SpO2 Averaging | Stable/Standard/Sensitive | Standard | SpO2-Averaging |
| Memory full | On/Off | On | Enables display of „memory-full“ message each start-up, if the device memory is full |
| Auto-off | On/Off | On | Turns device off after 2 minutes, if no measuring takes place |
| Factory settings | No/Yes | No | Reset device to factory settings |
| Exit |  |  | Exits menu |

System Information:

| Parameter | Displayed status (read only) | Description |
| --- | --- | --- |
| System failure code | 0x00 | System error code displayed as Hexadecimal value |
| Led damaged num | 0x00 | Number of LED errors displayed as Hexadecimal value |
| Work time | 0h 0min | Display of total working hours |
| Flash error num | 0 | Total number of Flash error |
| Module version | SWxxxx Vx.xx.xxxx | Version number of SpO2 Module |
| Module SN | xxxxWMxxxx | Serial number of SpO2 Module |
| LOT number | yyyy.mm | Year.Month of production |
| Exit |  | Exits menu |

# Maintenance

The OxyTrue® monitor is permanently factory calibrated. The maintenance and calibration-free SpO2 measurement technology ensures a robust measurement function throughout the lifetime of the monitor. However, to meet any local or specific customer requirements, a basic maintenance plan every 2 years is nevertheless highly recommended.

Recommended test steps are described in detail in the separate document OxyTrue Maintenance. All test steps involved in the maintenance check are to be documented in the separate document OxyTrue Maintenance Checklist. Both documents are available at the manufacturer upon request.

# Service and repairs

## General

Opening of the housing (particularly removal of screw-securing) is only allowed by authorised personnel trained by the manufacturer.

Warning:

After every manipulation inside the device (removal of screw-securing) the complete Service Protocol needs to be successfully completed before clinical use of the device to ensure proper operation. Contact the bluepoint service department (service@bluepoint-medical.com) for the latest Service Protocol Template.

Warning:

Remove batteries before each intervention to avoid damage to the circuits!

### Device disassembly

Remove the screw-lock with a screw driver. Remove all 6 screws and carefully remove the bottom housing from the front housing by lifting the backside slightly and shift left with low force to not damage the USB connector. Remove the white flat ribbon cable from the USB-PCB.

| Screw-lock |  | PCBA-screws |
| --- | --- | --- |
| Remove screw-lock and housing screws | Lift bottom housing, push slightly left. | Remove 6 PCBA screws to remove Main-PCB |

### Device assembly

Prior to inserting the screws, pay attention that no cables (especially the USB-PCB cable) are pinched by the screw holes. Do not tighten the screws with electric screwdriver but only with hand tools to avoid over winding. After the device is set back up together and bolted, insert the screw-lock in the upper right hole. Push gently inside without damaging the housing.

### OxyTrue® A components and HW compatibility

Main-PCB

Serial number

PCB-Screws

Battery contacts

Bottom housing

Front housing

SpO2 – Module

Flat ribbon cable

SpO2 sensor socket

USB-Connector

USB-Connection-cable

USB-PCB

Buffer capacity + Foam pad

Location of screw-lock

Device label

Back housing

Battery compartment

Battery compartment cover

Front housing

Display window

Keypad

Table 1: OxyTrue®A Hardware Compatibility List

| OxyTrue®A Main-PCB Version | SpO2 connector (Mounted on Main-PCB) | Compatible USB-PCB Version |
| --- | --- | --- |
| SMC20110215033(OLED) (discontinued) | 6 pin-Wire (REF: 2020522051) | USI20080807014 (discontinued) |
| SMC20131031033 (TFT) (discontinued) |  | USI20131031014 (discontinued) |
| SMC20141021033(TFT) (REF: 10020522001) | 6 pin-Flex (REF: 9020522001) | USI20141021014 (REF: 9020760001) |

### OxyTrue® S components and HW compatibility

Main-PCB

Serial number

PCB-Screws

Battery contacts

Bottom housing

Front housing

SpO2 – Module

SpO2 sensor socket

Flat ribbon cable

USB-PCB

Buffer capacity +

Foam pad

Location of 
screw-lock

Device label

Back housing

Battery compartment

Battery compartment cover

Front housing

Display window

Keypad

Table 2: OxyTrue® S Hardware Compatibility List

| OxyTrue® S Main-PCB Version | SpO2 connector (Mounted on Main-PCB) | Compatible USB-PCB Version |
| --- | --- | --- |
| SMC20110215033 (discontinued) | 6 pin-Wire (REF: 2020522051) | USI20080807014 (discontinued) |
| SMC20131031033 (discontinued) |  | USI20131031014 (discontinued) |
| SMC20141021033 (REF: 10020522051) | 6 pin-Flex (REF: 9020522001) | USI20141021014 (REF: 10020522050) |

## Replace Components

### Main-PCB

Warning:

Before replacing the Main-PCB, check the HW compatibility of the Main-PCB and the USB-PCB in Table 1 and 2 (section 4.1.3 and 4.1.4). The use of Main-PCB with incompatible USB-PCB may lead to malfunction and may affect patient safety!

Buffer battery for real-time-clock (RTC)

To replace the Main-PCB, open the housing as described in section 4.1.1. Disconnect all cables from the circuit board. Remove the SpO2 module and all 6 board screws. The Main-PCB can now be replaced.

| Back of Main-PCB with connectors and serial number label | Front of Main-PCB with display and buttons |
| --- | --- |

Before placing the new PCB, remove protective foil from the display. Remove dust at the display and display window with compressed air (clean and deionized, without oil). Place the new PCB into the upper housing and assure correct fitting.

Mount the PCB back on the housing with 6 screws. Insert the SpO2 module and connect all cables. Perform a function-test before fixing the housing parts with screws.

### USB-PCB

Warning:

Before replacing the Main-PCB, check the HW compatibility of the Main-PCB and the USB-PCB in Table 1 and 2 (section 4.1.3 and 4.1.4). The use of Main-PCB with incompatible USB-PCB may lead to malfunction and may affect patient safety!

Remove the defective USB-PCB by opening the housing as described in section 4.1.1. Disconnect all cables and remove the battery leads from the bottom housing with a soldering iron. Remove the three USB-PCB screws, take out the PCB and remove the buzzer cables at the USB-PCB with a soldering iron.

| Soldering of battery leads | Soldering of buzzer leads |
| --- | --- |

Take the new USB-PCB (without USB-PCB cable for model OxyTrue® S) and solder the buzzer cables to the USB-PCB. Place the USB-PCB on the guiding pins at the bottom housing without pinching the buzzer cable. Screw the PCB to the bottom housing and solder the battery leads to the terminals at the back housing. Perform a function-test before fixing the housing parts with screws.

### USB-PCB cable

Replace the USB-PCB cable (OxyTrue® A only) if no USB connection can be established. Open the housing (section 4.1.1) and disconnect the cable from both plugs, replace the cable with a new cable (REF: 2020522052). Both cable ends should be inserted in the plug (see image below).

Warning:

Pay attention that no cable is clamped inside the screw-holes while mounting the housing back together.

| Correctly fitted cable | Position cable between USB connector and screw shaft to prevent clamping at screw hole. |
| --- | --- |

### Flat ribbon cable

Replace the Flat ribbon cable if it is mechanically damaged or if the power supply of the device shows malfunction.

Open the housing (section 4.1.1) and disconnect the cable from both plugs, replace the cable with a new cable (REF: 2020522054). Both cable ends should be inserted in the plug (see image below). Fold the cable to the direction of the USB plug while closing the device (not in direction of the SpO2 module) to prevent clamping.

Fold cable into the direction of the USB plug to prevent clamping while closing the device

### SpO2 Module

If the module is defective or not connected correctly, the message “Device defective!” will be displayed. Open the housing (section 4.1.1) to verify correct fitting or replacement of the module.

First check correct fitting to avoid unnecessary replacements. The SpO2 module must be connected to the Main-PCB in alignment with the white frame marking on the Main-PCB. All pins must be connected to the socket (no pin may be visible).

| Module aligns over the white frame on the Main-PCB | Both plugs must mount completely on the sockets. |
| --- | --- |

For repositioning remove the module and check the soldering of the socket at the module. Replace the module if the socket solder points are damaged/loose. In case of no damage at the socket, reposition the module correctly and perform a functional check. Replace the SpO2 module if the error persists.

After replacing the module, test the SpO2 measurement function. Replace the Main-PCB if the malfunction continues (chapter 4.2.1).

### SpO2 sensor socket

If the device is showing “Sensor fault”, or the measurement is interrupted from time to time, the SpO2 sensor socket may be faulty.

If a sensor failure can be excluded, the SpO2 sensor socket has to be replaced. To exclude a sensor failure, perform a measurement with a known working sensor. If the failure persists, replace the SpO2 sensor socket.

Open the housing (section 4.1.1) and disconnect the flat ribbon cable. The device can be opened completely. Remove the SpO2 sensor socket and replace it with a new one.

Select the SpO2 sensor socket depending on the HW version of the main PCB. See Table 1 and 2 (section 4.1.3 and 4.1.4) to order the correct spare part.

| 6 pin-Wire (REF: 2020522051) | 6 pin-Flex (REF: 9020522001) |
| --- | --- |

NOTE: The dark square at the flex of the SpO2 sensor socket must face the SpO2 module, else the message “Sensor fault” will continue.

The nose at the SpO2 sensor socket has to point to the display side after inserting it into the front housing.

| SpO2 module |  |
| --- | --- |
| The dark square at the flex must face the SpO2 module. | The nose at the SpO2 sensor socket has to point to the display side after inserting. |

After replacing the SpO2 sensor socket, test the SpO2 measurement function by connecting a sensor to the finger. Wiggle the sensor connector in the plug and check that the plethysmogram is not influenced by the movement of the connector. The measurement has to be continuous with no error message.

| The Plethysmogram may not be influenced by wiggling the plug. | No error messages may occur while wiggling the plug. |
| --- | --- |

### Keypad

Before replacing the keypad, pay attention to order the correct version (OxyTrue® A/ S have different layouts).

Open the housing as described in section 4.1.1 and remove the Main-PCB. Be careful not to touch the inside of the display window. Place a scalpel or similar tool carefully under the lower left edge of the keypad. Lift the keypad and remove it carefully from the top housing (with finger or pliers).

Glue remains can be removed with fingers or suitable tool without scratching the front housing. Clean the surface of the housing with 70% isopropanol.

Remove the supporting foil from the new keypad and place it on the lower rim of the front housing. Stick it carefully around the display window and fix it by pressing with the finger.

| Attach tool and remove the keypad. | Remove the supporting foil | Fix the new keypad to the housing |
| --- | --- | --- |

Remove dust at the display and display window with compressed air (clean and deionized, without oil). Place the Main-PCB into the upper housing and assure correct fitting. Close the device according to section 4.1.2.

NOTE: It is possible to replace the keypad without opening the device.

### Display window

Open the housing as described in section 4.1.1 and remove the Main-PCB.

Place a flat screw driver or similar tool between the display window and the housing to remove the display window. Remove dirt and glue remains from the housing using 70% isopropanol.

Remove the backside protecting foil from the new display window. Do not touch or damage the inner surface of the display. Fingerprints at the inner surface of the display window cannot be removed.

| Remove the display window with a flat screw driver | New display window |
| --- | --- |

NOTE: The display should now be opaque. If not, remove the middle protection foil carefully without removing the outer silvery foil. Insert new window in an appropriate way. Remove dust at the display and display window with compressed air (clean and deionized, without oil). Place the Main-PCB into the upper housing and assure correct fitting. Close the device according to section 4.1.2.

### Housing

Bottom housing

A spare serial number label with transparent protection sheet is required during replacement of the back housing. Order the label by contacting the bluepoint service department (service@bluepoint-medical.com). State your device name and serial number.

NOTE: The OxyTrue® A and S have differed label.

Replace the bottom housing by open the device and removing the USB-PCB (refer to section 4.2.2).

Place the label on the housing. Ensure that the buzzer hole is not covered by the label to avoid reduction of alarm sound volume. Close the housing according to section 4.1.2. Connect a new screw-lock.

Serial number Label with protective sheet

Screw-lock

Device Label

NOTE:

Ensure that the buzzer hole is not covered by the label to avoid reduction of alarm sound volume.

Front housing

Replace the front housing by open the device and removing the Main-PCB (refer to section 4.2.1).

If no window or keypad is fitted, place keypad and window according to section 4.2.7 and 4.2.8.

Close the housing according to section 4.1.2.

# Troubleshooting Guide

| ID | Failure mode | Possible causes | Corrective action | see Chapter |
| --- | --- | --- | --- | --- |
| #1 | No display function (stripes, dark) | display defective, control of display defective | Replace Main-PCB | 4.2.1 |
| #2 | Keys without function | Main-PCB or buttons on Main-PCB defective | Replace Main-PCB | 4.2.1 |
| #3 | No USB connection | PC to USB cable defective, PCB-connector-cable faulty, USB-Connector on USB-PCB or Main-PCB defective | Replace cable(s) or Main-PCB | 4.2.3 / 4.2.1 |
| #4 | Data download not possible despite recognized USB connection | Incompatibility of PC configuration, USB-PCB cable faulty, USB-Connector on USB-PCB or Main-PCB defective | Check PC configuration and USB drivers. Replace cable(s) or Main-PCB | 4.2.3 / 4.2.1 |
| #5 | Keypad mechanically damaged | Wear of Keypad due to mechanical stress | Replace Keypad | 4.2.7 |
| #6 | Clock is set to zero after batteries were changed | Buffer battery on Main-PCB flat or Main-PCB defective | Replace buffer battery on main PCB if voltage below 3.1V. Replace Main-PCB | 4.2.1 |
| #7 | Display window defective (scratched, broken) | Wear of display window due to mechanical stress (fall damage) | Replace display window | 4.2.8 |
| #8 | Device displays „Sensor fault“ | SpO2 sensor defective or SpO2 sensor socket faulty | Replace SpO2 sensor or SpO2 sensor socket | 4.2.6 |
| #9 | Interrupted measurement |  |  |  |
| #10 | Device displays „Device defective!“ | SpO2 -Module faulty, wrong/defective positioned SpO2 Module or Main-PCB defective | Correct/replace SpO2 -Module or replace Main-PCB | 4.2.5 / 4.2.1 |
| #11 | Strong fluctuation of measured values | - Patient has low perfusion / Motion artefacts - Sensor faulty, SpO2 Module or sensor socket defective | - Warm hands of patient, ensure no motion artefacts - Replace sensor, replace SpO2 Module | 4.2.6 / 4.2.5 |
| #12 | ERROR code 01 | Refer to ID #10 (SpO2 -Module faulty, wrong/defective positioned |  |  |
| #13 | ERROR code 02 | Error in flash for stored data | Delete all data. If the error continues, replace Main-PCB | 4.2.1 |
| #14 | ERROR code 03 | Buffer Battery for Real Time Clock flat (should be 3.1V) | Set time in the main menu, replace main PCB if the error continues. | 4.2.1 |

# Spare parts

All available spare parts are listed in the table below with separate part numbers (REF) for OxyTrue®A/ S.

Fig. 1: Exploded View – OxyTrue® A

For hardware compatibility of parts see section 4.1.3 and 4.1.4.

| Part number (REF) |  | Description | Identifier (Fig.1) |
| --- | --- | --- | --- |
| OxyTrue®A REF: 1020112001 | OxyTrue®S REF: 2020112011 |  |  |
| 5020522001 | 5020522002 | Front housing without keypad and display window. (white for OxyTrue® A, grey for OxyTrue ®S) | 3, 4 |
| 2020522009 | 4020522006 | Bottom housing with battery cover and buzzer. (dark blue for OxyTrue® A, light blue for OxyTrue ®S) | 6, 7, 8, 9, 13, 14, 15 |
| 2600522013 |  | Screw set housing (2 long, 4 short screws) | n/a |
| 2020522056 | 4020522004 | Battery cover (dark blue for OxyTrue® A, light blue for OxyTrue ®S) | 9 |
| 2020522014 |  | Display window | 2 |
| 2020522015 | 4020522000 | Keypad (different layout for OxyTrue® A and S) | 1 |
| 2020522051 |  | 6 pin-Wire SpO2 sensor socket (mounted on Main-PCB version SMC20110215033 and SMC20131031033) | 11 |
| 9020522001 |  | 6 pin-Flex SpO2 sensor socket (mounted on Main-PCB version SMC20141021033) | 11 |
| 2020522054 |  | Flat ribbon cable (power connection at Main-PCB to USB-PCB) | n/a |
| 2020522052 | n/a | USB-PCB cable (USB connection at Main-PCB to USB-PCB | n/a |
| 9020760001 | 10020522050 | USB-PCB version USI20141021014 | 12 |
| 10020522001 | 10020522051 | Main-PCB version SMC20141021033(with TFT display) | 5 |
| 5600522005 |  | SpO2 module – ChipOx | n/a |

# Document History

| Rev. | Date | Issued | Reviewed | Description of changes |
| --- | --- | --- | --- | --- |
| 0 | 18.07.2008 | HB | TB | initial version Manual for Service Menu Options, no repair instructions |
| 1 | 07.11.2011 | SB | SB | add (1.): Parts overview add (3.): Repair instructions (open / close housing, FW update, repair of: keypad, display window, housing, Main-PCB, USB-PCB, Usb-PCB cable, SpO2 module, minimed socket) add (4.): Spare part overview former: 6.0_service_manual-oxytrue_rev1.1.docx |
| 2 | 13.06.2012 | SB | SH | add (3.7): exchange of OLED add (4.0): OLED part number (REF) former: 6.0_service_manual-oxytrue_rev1.2.docx |
| 3 | 27.11.2017 | HF | SH, JPB | Change section order and update syntax. Update images to new PCB versions. Update and correct part numbers. Add sections 4.1.3, 4.1.4, remove former section 3.2. |