## Rev. A0
1. Select a suitable site for the sensor. The patient`s Index ﬁnger is the preferred location. Alternative sites recommended are the thumb, large toe,
or smaller ﬁnger.
2. Remove the adhesive backing from the sensor. Place the patients digit nail-side up directly over the detector.
3. With the sensor positioning line running across the nail, wrap the bottom adhesive around the digit. Do not cover the nail.
4. Fold the sensor’s top over the digit and align the light source directly opposite the detector .
5. Wrap the adhesive arround the digit and secure with tape if necessary.
6. Connect the sensor cable to the monitor or to the patient cable and verify proper operation as described in the instrument's operator manual.
Indications |  |  |  | Warnings and precautions
The sensor is indicated for use for continuous non-invasive or spot |  |  | 1. | This sensor is for use only with compatible monitors,instruments, or oximetry
monitoring of functional oxygen saturation of arterial hemoglobin (SpO2) |  |  |  | modules. Use of this sensor with instruments other than compatibles
and pulse rate for patients in hospitals, hospital-type facilities, mobile. |  |  |  | may result in improper performance.
|  |  | 2. | Verify proper operation before use.
The sensor is indicated for single patient use only. |  |  | 3. | Operations may be affected in the presence of strong electromagnetic
sources such as electrosurgery equipment.
Patient weight |  |  | 4. | The product is not recommended for use in the presence of imaging
Recommended for use on patients weighing between 10kg and 50kg / |  |  |  | equipment such as Magnetic Resonance Imaging (MRI) and Computed
22lbs - 110lbs. |  |  |  | Tomography (CT) devices, etc.
|  |  | 5. | Do not autoclave or immerse in liquid of any kind. Do not sterilize with ETO.
Accuracy |  |  | 6. | Do not stretch the adhesive tape too tighten to assure that the blood-
1. SpO2 70% - 100% SpO2 ± 2 digits |  |  |  | circulation is not restricted.
< 70% SpO2 : unspecified |  |  | 7. | If the sensor is damaged in any way, discontinue use immediately.
|  |  | 8. | Operation may be affected in the presence of high ambient light (also
Verified by clinical tests, where the measured values of the sensors were |  |  |  | because of IR and UV light). Shield the sensor area with a surgical
compared with those of the arterial co-oximetry in adult subjects over the |  |  |  | towel if necessary.
specified SpO2 range. |  |  | 9. | Check and reposition the sensor to an alternate location every 4 hours.
|  |  | 10. | Intravascular dyes may interfere with performance of the sensor and
2. Pulse Rate 30 - 250 BPM : ±3 BPM |  |  |  | the monitor which may cause inaccurate measurements.
|  |  | 11. | The performance of the sensor and the monitor may be compromised
by excessive motion. To decrease this motion, make a loop with the
Storage temperature - 20°C - + 70°C |  |  |  | sensor cable and secure it to the patient using tape.
Operating temperature + 5°C - + 40°C |  |  | 12. | Fingernail polish or an artificial fingernail can cause inaccurate SpO2
readings.
|  |  | 13. | Refer to the monitor’s operations manual for additional cautions and
Symbols and indications |  |  |  | warnings.
|  |  | 14. | Significant levels of dysfunctional hemoglobin’s, such as Carboxyhemoglobin
| Do not re-use. Single use product |  |  | or Methemoglobin will affect the accuracy of the oxygen measurement.
|  |  | 15. | Application of the sensor to edematous or fragile tissue must be avoided.
| Latex FREE |  | 16. | Venous pulsations that could potentially lead to inaccurate saturation
measurements may be caused when the sensor is wrapped too tightly
| Non-sterile product |  |  | or supplemental tape is applied.
Sensatronic GmbH
| Dispose according to the local |  |  |  | Am Ring 9
| legislation on W.E.E.E. |  |  |  | 23970 Wismar
Germany
Consult instructions for use
Shenzhen Envisen Industry Co. Ltd.
| Keep away from sunlight |  |  |  | 2nd Floor, Block 1, 40 Jianlong Street,
Baoan Community, Heng Gang Town,
Longgang District Shenzhen , China
Catalogue number
Disposable SpO2 Sensor |  | pediatric
Model: | DP-2231 | Nonin compatible
Type: | 7-PP-VM (Plaster)
Manufacturer
European conformity
Use-by date
IPX2
IP Classification indicating
Temperature limit
Keep dry
For prescription use only
Lot number