[image: image1.jpg]


VIAMED

       INFANT RESUSCITATION

CABINET ASSEMBLY – Part Number 0310002

VM3/COP/51.01

Date:
Revision Date:5-Sep-08
Issue: 1


[image: image2.png]Adjustable distance from walll
maximun : 1100 mm

4———— 550 mm ~——

90 mm‘
280mm. »

Between 800mm & 1000mMm

") Mattress Height
820 mm—————=p

Approximately 1 Metre
Chosen by User

Floor





Equipment required:
Spirit Level, Electric Hand Drill, ?? drill bits, 1 Set crosshead screwdrivers, Centre Punch, Hammer, Isopropyl Alcohol, Kitchen Tissue.

Parts List

Quantity
Description
Part Number

1
Cabinet. 
(As per order)

1
Ceratherm 600-2 Infant Radiant Warmer.
(As per order)

1
Plastic Bed
0320019

1
Mattress
(As per order)

?
Rubber Feet






Accessories List

Quantity
Description
Part Number

1
Tom Thumb Resuscitation Unit.
(As per order)

1
Suction Controller.
(As per order)

1
APGAR Timer.
(As per order)

3
Lin-Bins.
(As per order)

1
Hose. (1m or 3m)
(As per order)









Cabinet Assembly

1. Taking care not to scratch the surface paint, remove the cabinet from the protective wrapping and place on the workbench. Examine the whole cabinet for scratches, blemishes and aesthetic faults: such as paint flakes, discolouring etc.

2. Open the clasp and drop down the front door. Remove the internal parts (Medirails). Examine the inside of the cabinet for scratches, blemishes and other aesthetic faults etc. Check that the drop down door is laying at a true “Right Angle” to the cabinet.














3. Check all the hinges and especially the door clasp. There must be no play in these items, nor must they be over-tight. The door when closed must not be seen to be buckling, or hanging ajar.

4. Arrange the products to be affixed, on the bench, in the order of assembly.

5. Using the screws and nylon locknuts supplied, affix the Medirails to the back wall of the cabinet. One to the top set of holes, one to the middles set of holes and one to the lower set of holes.

90º





Wall


Cabinet





Cabinet Door












































Page 3 of 3

