CC Issue #13921 Microstim DB3 mode switching query

Original Issue# 13918
Created: 10 Mar 2008
Status Still Outstanding

Copy Issue Main Issue #13918
Added by steve nixon
Reported by a customer: If you flick from one \`paired\` selection to the other, the original programme remains active i.e. from PTC to 1Hz.

This is correct, but this has always been the case. You need to switch very fast and in practice this would not be done. The anaesthetist using the unit would soon spot this. In any case with TOF and DBS not an issue since programme stops after one cycle. With PTC and 1Hz not a real problem since PTC goes to 1Hz if you keep finger depressed. 

Suggest add note to next revision of the manual:

``When switching between two modes of operation first release the switch before activating the next mode``.

