VM 2101

Comments – Prototype II

1) Case matt finish is good, original was gloss finish.

2) Revised Viamed logo is good.

3) Support pillars thickened – good.

4) Lateral movement much improved.

5) Moulding sink marks have been resolved – good.

6) Visible screen size has been enlarged by about 1mm in both dimensions -good.

7) Display brightness improved – good.

8) Blue graphics have been lightened – good.

9) Battery compartment cover much improved, still slightly loose when closed. Bernd says this is being resolved.

10) Battery compartment, the wires were visible on the first prototype at each end of the battery terminals. This has been resolved with revised moulding so that no wires visible at all – good.

11) Battery compartment, no polarity indication. Bernd says this is being resolved.

12) Wiring cover between the two case sections is good, suggest making say black instead of light grey. Bernd says this is being resolved to colour match. Existing sample is just for EMC testing purposes.

13) Brightness adjustment added to software, 6 stages – good. No indication of level, Bernd says this is being resolved with use of a bar graph and ISO symbol for brightness.

14) Bernd stated that the ‘Finger Out’ message is to be replaced with ISO symbol to avoid translation issues. If this is so, do we also change BPM to a symbol?

15) Finger pads, colour is good but does not look uniform. May be just dirty or needs extra dye adding to the silicone?

16) Low perfusion capability will be an advantage. Need testing data to substantiate.

17) 16 bit A-D will also be god for marketing. I assume that the original design utilized 8-bit A-D processing?

30th May 2008
