DVM Calibration record

DVM Calibration Index Number
CE006

DVM Bar code
ID101990

Equipment Used
Time 1030 MicroCal

Serial number Test Equipment
6839k6

Calibration Certificate Number
1098426

Test Equipment Calibration Index Number
CE076 (Barcode 102060)




Micro Cal @
DVM Reading

0mV
0.0

5mV
5.0

10mV
9.9

15mV
14.9

20mV
19.9

25mV
24.9

30mV
29.9

35mV
34.8

40mv
39.8

Pass / Fail
Pass

Tested By
Derek Lamb

Tested On
02/06/2008

Signed


