DVM Calibration record

DVM Calibration Index Number
CE110

DVM Bar code
ID102094

Equipment Used
Time 1030 MicroCal

Serial number Test Equipment
6839k6

Calibration Certificate Number
1098426

Test Equipment Calibration Index Number
CE076 (Barcode 102060)




Micro Cal @
DVM Reading

0mV
0.0

5mV
5.0

10mV
10.0

15mV
15.0

20mV
20.0

25mV
25.0

30mV
30.0

35mV
35.0

40mv
40.0

Pass / Fail
Pass

Tested By
Derek Lamb

Tested On
02/06/2008

Signed


