## Slide 1
MDSAP Overview
Damon Williams
Business Development Manager - BSI

## Slide 2: What is

## Slide 3: MDSAP Update
Agenda
What is MDSAP
How MDSAP Works
How does MDSAP fit with other certifications
MDSAP Status

## Slide 4: Medical Device Single Audit Program (MDSAP)
Result of one of the 6 Working Groups created by the International Medical Device Regulatory Forum (IMDRF)
Global approach to auditing and monitoring the manufacturing of medical devices to ensure safe medical devices
An international coalition to quickly pilot the program
Objective:
  to jointly leverage regulatory resources to manage an efficient, effective, and sustainable single audit program focused on the oversight of medical device manufacturers.
History
Notes: WHAT IS MDSAP?
The MDSAP Work Group was created at the inaugural IMDRF meeting Sep 2012.
An international coalition quickly agreed to pilot the program from Jan 2014 to Dec 2016.
4 regulatory bodies: Australia, Brazil, Canada, USA initially, with Japan as observer. June 2015 Japan officially part of coalition

## Slide 5: MDSAP
Develop, manage, and oversee a single audit program that will allow a single regulatory audit to satisfy the needs of multiple regulatory jurisdictions
To promote greater alignment of regulatory approaches and technical requirements
To promote consistency, predictability, and transparency of regulatory programs
Program Objectives

## Slide 6: MSDAP
No additional requirements for manufacturers
Single audit optimizes time and resources
Routine audits are scheduled/planned with AO
Expected to improve predictability
Expected to add additional Regulatory Authorities
Manufacturer Benefits

## Slide 7: MDSAP Program Status
Use of outputs of MDSAP audits
Australia | Brazil | Canada | Japan | USA
Use as part of evidence to assess compliance with MD market authorization requirements | Input for ANVISA’s premarket and post-market assessment procedures | Concurrent with CMDCAS until ends in late Dec 2018 | Report might be utilized for a desk review for class 2,3,4 in lieu of a premarket inspection performed by PMDA or registered certification bodies in Japan | Substitute for Routine Inspections only. Not for PMA, “For Cause” or “Compliance Follow-up”
 | Audits in lieu of ANVISA inspection to grant GMP certs for class 3,4 | Use of certificate for obtaining/maintaining a Class 2,3,4 device license | report might also be utilized for periodic post market inspections | Report review with scrutiny on significance of findings
 | For renewal of ANVISA’s GMP certs bi-annually | From January 2019, Health Canada will only accept MDSAP certificates | Reports will be used in review of on-site inspection for eligible sites so as to obtain a QMS certificate | May use Warning Letters if conclusion of imminent/unreasonable risk to public health

## Slide 8: How MDSAP works

## Slide 9: MDSAP Structure
Regulatory Authorities
Assessments
Auditing Organizations
Audits NCs/CAPs
Manufacturers
NCs/CAPs
Reports
Notes: In a nutshell, the RA’s assess and qualify the AO’s,
Qualified AO’s audit the Manufacturers and send the reports to the AO’s.
Procedures have been implemented to ensure consistency across the RA’s and the Assessments that they conduct of the AO’s, as well as the AO’s, and the Audits that they conduct of the Manufacturers.

## Slide 10: MDSAP
Three Year Audit Cycle
  Initial Audit (Stage One & Stage Two)
  Surveillance Audits (Years 1 and 2)
  Re-audit (Recertification Audit)
    Note that not all Regulatory Authorities require “certificate”
Other Possible Audits
  Special Audits
    changes, nonconformances, suppliers, post-market issue follow-up
  Audits by Regulatory Authorities
  Unannounced Audits (to close major nonconformances)
Audit Cycle
Notes: Now for a deeper dive into the audits themselves.

## Slide 11: MDSAP
ISO 13485 plus applicable Country-specific requirements
A separate report is required per site.
To recommend certification to MDSAP all applicable processes and jurisdictions must be audited.
There is no sampling of design and manufacturing sites permitted in the MDSAP program.
All RA’s in the program get copies of submitted reports
  Not subject to US FOI Act
Program Requirements
Notes: Requirements to which manufacturers will be audited to consists of two primary parts:
ISO 13485 requirements, and Country-specific requirements
These are clearly identified in the Audit Model document
If a manufacturer is shipping product to a MDSAP jurisdiction, country-specific requirements WILL apply

## Slide 12: Regulations in addition to ISO 13485
 | Requirements
 | Therapeutic Goods Act 1989 Therapeutic Goods (Medical Devices) Regulations 2002
 | ANVISA Pre-Market Approval RDC 185/2001 ANVISA Good Manufacturing Practices RDC 16/2013 ANVISA GMP Certification – Requirement for Product Registration RDC 25/2009 ANVISA PMS RDC 67/2009 and RDC 23/2011
 | Food and Drugs Act R.S.C., 1985, c. F-27 CMDR SOR-98-282
 | Quality System Regulation 21 CFR 820, Medical Device Reporting 21 CFR 803, Reports of Corrections & Removals 21 CF 806, Registration & Listing 21 CFR 807 subparts A to D, Device Tracking 21 CFR 821
 | MHLW Ministerial Ordinance No. 169

## Slide 13: MDSAP Audit Sequence
Notes: Here is a graphic representation of the sequence. Would basically follow the primary processes, but linkages to the supporting processes could take place as the audit progresses. Parts of Purchasing could be looked at during any of the primary processes. Risk Management would be expected to underlie all processes.
Follows the process approach, top down
Four Primary processes
  Management
  Measurement, Analysis and Improvement
  Design and Development
  Production and Service Controls
Three Supporting Processes
  Device Marketing Authorization and Facility Registration
  Medical Device Adverse Events and Advisory Notices Reporting
  Purchasing

## Slide 14: MDSAP Audit Process Sequence and Estimated Durations
MDSAP Process | MDSAP Tasks per Process | Minutes per Audit Task
Management | 11 | 36.0
Device Marketing Authorization &Facility Registration (DMA&FR) | 3 | 35.0
Measurement Analysis & Improvement (MA&I) | 16 | 38.0
MD Adverse Events & Advisory Notice Reporting (MDAE&ANR) | 2 | 38.0
Design & Development (D&D) | 17 | 21.0
Production & Servicing Controls (P&SC) | 29 | 44.0
Purchasing | 16 | 15.0

## Slide 15: MDSAP Audit Time Calculations SAMPLE
hh:mm
* D&D and P&SC can be split between Surv 1&2, Purchasing to follow based on device trail.
 | Initial Cert. | Surv 1* | Surv 2* | Recert.
Management | 6:36 | X | X | 6:36
DMA&FR | 1:45 | X | X | 1:45
MA&I | 10:08 | X | X | 10:08
MDAE&ANR | 1:16 | X | X | 1:16
D&D | 5:57 | x | X | 5:57
P&SC | 21:16 | X | x | 21:16
Purchasing | 4:00 | x | x | 4:00
On-site Total | 50:58 (6.5d) | ~70% of cert. | ~70% of cert. | 50:58 (6.5d)

## Slide 16: MDSAP Audit Program
Considerations for audit time adjustments
Past Nonconformance review adds time to MA&I
D&D tasks adjustments:
  No design then only task 1 & 16
  No active role with new device design or devices prior to regulatory design requirements tasks limited to 1,4 & 13-16
  Devices containing SW may result in duplicate D&D tasks
P&SC task adjustments
  Multiple processes will add time per process
  Limited processes and no changes since last audit may reduce time
Purchasing time will be dependent on number of critical suppliers and will include intra-company entities, where applicable

## Slide 17: MDSAP
Uses GHTF Document SG3/N19:2012 - Nonconformity Grading System for Regulatory Purposes and Information Exchange
Definition of nonconformity unchanged – non-fulfillment of requirement
Creates a quantitative grading system
Nonconformity Grading
Notes: Nonconformity grading differs from minor/major used in ISO audits

## Slide 18: MDSAP Nonconformity - Initial Grading Matrix
Impact
Occurrence
QMS Impact on safety & performance | Direct Clauses 6.4 thru 8.5 | 3
 | Indirect Clauses 4.1 thru 6.3 | 1
Occurrence in same sub-clause (X.X.X | Repeat during last 2 QMS audits | +1
 | First | 0

## Slide 19: MDSAP Nonconformity - Escalation Rules
Escalation
Matrix
Absence of documented process or procedure
Release of a nonconforming medical device
Absence of Process or Procedure | +1
Led to Nonconforming devices on market | +1
Escalation Criteria | 

## Slide 20: MDSAP Nonconformity Grading - Final
Maximum grade is a 5.
QMS Impact | Direct | 3 | 4 | Absence of Process or Procedure | +1
 | Indirect | 1 | 2 | Led to Nonconforming devices on market | +1
 |  | First | Repeat |  | 
 |  | Occurrence |  | Escalation Criteria | 

## Slide 21: MDSAP Program Update
Manufacturer’s next steps
Assess markets for the sale of your product now and in the future
If you have are due for recertification to CMDCAS and intend to market in Canada you should consider transition to MDSAP when you are up for recertification.
  Those who have already recertified in 2016 you will need to become certified to MDSAP before your next recertification comes up.
Coordination of MDSAP and ISO 13485:2016 is crucial in your planning process.
Contract for your MDSAP audit - ASAP time is of the essence!!!
Use On-Line survey to provide feedback to the regulators about your experience with the MDSAP program. http://www.fda.gov/MedicalDevices/InternationalPrograms/MDSAPPilot/default.htm

## Slide 22: MDSAP Program Update
QUESTIONS ???