QC 21 Non Conformance Report
	Date
	24/09/2019

	Issue id
unique identifier 
	155083

	BSI Ref (if applicable) 
unique identifier 
	1826110-201909-N2 

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	 

	Investigation By:
Person responsible 
	Derek Lamb

	Investigation Issue id
(if applicable)
Root Cause Analysis 
	The grease and sealant used in the workshop for servicing was not reviewed regularly to ensure it had not gone past its manufacturer recommended life span.
With them not having a use by date, we missed reviewing the product and checking their data sheets for a manufacturers life span of the product.  

	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	

	Time Scale for Corrective Action
Time for completion of all identified actions 
	Immediate Action to be completed by 19th October 2019
Corrective action to be completed by 19th October 2019


	Corrective Action:
	Immediate Action
A questionnaire will be sent to all relevant staff to list any sealant, glues, greases, sprays and tapes that are used in production, servicing and repairs for Viamed and VST. To check their data sheets and make sure none are over their use by date or manufacturers shelf life. 
We are going to add regular issue to check the Data Sheets on products used, for a life span, and dispose of them when they reach what the manufacturer states is the shelf life. 
A question will be added to Audit 11 Repairs, Servicing and Returns to review this. 
Corrective Action
All all the relevant procedures will be updated to state that a check needs to be carried out on all products used in production and servicing, to make sure they are within their use by dates or manufacturers shelf life. 
Moving forwards will have regular reviews and issues sent to make sure that all new and existing products used in production or servicing, are looked at regularly and disposed of when they are past the use by date or shelf life.
All current and new relevant procedures will include this process. To make sure we do not miss this again. 


	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


