QC 21 Non Conformance Report
	Date
	24/09/2019

	Issue id
unique identifier 
	155083

	BSI Ref (if applicable) 
unique identifier 
	1826110-201909-N2 

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	The product preservation is not fully effective as no process was
seen in place to ensure that the chemicals (grease and sealant
used respectively for the lubrication and to lock the regulator
screw of the Tom Thumb infant resuscitator) were used within
their lifetime.


	Investigation By:
Person responsible 
	Derek Lamb

	Investigation Issue id
(if applicable)
Root Cause Analysis 
	The grease and sealant used in the workshop for servicing was not reviewed regularly to ensure it had not gone past its manufacturer recommended life span.
With the original manufacturer not having a use by date or expiry date on the product. We missed reviewing these products and we did not have a process in place, to check the manufacturers data sheets for a life span of the product.  


	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	Issue ID154977
Issue ID154978
Task ID7940 
Task ID7944


	Time Scale for Corrective Action
Time for completion of all identified actions 
	Immediate Action to be completed by 19th October 2019
Corrective action to be completed by 19th October 2019


	Corrective Action:
	Immediate Action
A questionnaire will be sent to all relevant staff to list any sealant, glues, greases, sprays and tapes that are used in production, servicing and repairs for Viamed and VST. To check their data sheets and make sure none are over their use by date or manufacturers shelf life. 
A new process will be added in to Intrastats, which will generate the requirements of all the processes eg. Risk Assessments, Measurable Objectives and Training Requirements. This will generate Rolling Task, Rolling Issue and Rolling Audit Task.
A Rolling Task will be added to check the Data Sheets on products used, for a life span. This Task will also included a statement to tell staff to dispose of any that have reached what the manufacturer states is the shelf life. 
The Rolling Audit will Ensure the task is completed in a timely manner each year.
This process is a new addition to the system and does not affect any existing processes.
A question will be added to Audit 11 Repairs, Servicing and Returns to review this Process. 
Corrective Action
All all the relevant procedures will be updated to state that a check needs to be carried out on all products used in production and servicing, to make sure they are within their use by dates or manufacturers shelf life. 
Moving forwards will have regular reviews and issues sent to make sure that all new and existing products used in production or servicing, are looked at regularly and disposed of when they are past the use by date or shelf life.
All current and new relevant procedures will include this process. To make sure we do not miss this again. 


	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


