QC 21 Non Conformance Report
	Date
	24/09/2019

	Issue id
unique identifier 
	155084

	BSI Ref (if applicable) 
unique identifier 
	1826110-201909-N3 

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	 

	Investigation By:
Person responsible 
	Derek Lamb

	Investigation Issue id
(if applicable)
Root Cause Analysis 
	We did not have suitable QA / test procedure in place for the testing of the MD300 Oxywatch.


	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	

	Time Scale for Corrective Action
Time for completion of all identified actions 
	Corrective action to be completed by 
19th October 2019

	Corrective Action:
	We will correct the relevant procedure and add issues to make sure, if procedures are missing or not suitable, they are reported back to management.
Review all procedures of incoming stock so we can check their suitability. We will put issues in place to remind staff and review regularly.
We will put a system in to Intrastats to monitor incoming goods and their procedures. To check if they are present and when they were last reviewed. 
Add rolling tasks and audits to make sure these are reviewed regularly. 


	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


