QC 21 Non Conformance Report
	Date
	24/09/2019

	Issue id
unique identifier 
	155084

	BSI Ref (if applicable) 
unique identifier 
	1826110-201909-N3 

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	The verification of purchased product is not fully effective as procedure for the incoming inspections was seen requiring verification activities to be perfromed and verified against
acceptable limits, however no limits were seen defined.
 

	Investigation By:
Person responsible 
	Derek Lamb

	Investigation Issue id
(if applicable)
Root Cause Analysis 
	We did not have suitable QA / test procedure, in place, for the testing of the MD300 Oxywatch.
The product in question is an OEM product, carrying its own CE mark. In the past we had a few Dead on Arrival failures with this companies products. So a Go – No Go QA test was introduced, so we could deal with failures immediately and not have the end users be affected. This QA procedure incorrectly implied we should be taking or at least ensuring suitable readings without a means of accurate measurement.


	Corrective Action By:
Person responsible 

	Derek Lamb

	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	Process ID 7942
Issue ID 156113
Issue ID 156100 
Issue ID 155531
Issue ID 154997
Issue ID 154996
Issue ID 154995
Issue ID 154994
Issue ID 154994
Issue ID 154993
Issue ID 154992
Issue ID 154991
Issue ID 154644

	Time Scale for Corrective Action
Time for completion of all identified actions 
	Immediate Action to be completed by 19th October 2019
Corrective action to be completed by 19th October 2019


	Corrective Action:
	Immediate Action
We will correct the relevant procedure and add issues to make sure, if procedures are missing or not suitable, they are reported back to management.
We will discuss with all relevant staff, the importance of having the correct, up to date procedure for each incoming product. So they can verify these, whenever they process stock coming in to the building. 
Any procedures that need to be updated will be done immediately. 
Corrective Action
Review all procedures of incoming stock so we can check their suitability. We will put issues in place to remind staff and review regularly.
A new process will be added in to Intrastats, which will generate the requirements of all the processes eg. Risk Assessments, Measurable Objectives and Training Requirements. This will generate Rolling Task, Rolling Issue and Rolling Audit Task.
We will put a system in to Intrastats to monitor incoming goods and their procedures. To check if they are present and when they were last reviewed. 
The rolling tasks and audits will ensure these systems and procedures are reviewed regularly. 


	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


