QC 21 Non Conformance Report
	Date
	24/09/2019

	Issue id
unique identifier 
	155085

	BSI Ref (if applicable) 
unique identifier 
	1826110-201909-N4

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	The risk management process is not fully effective as no linkage was provided to the verification activities of the risk control
measures and the assessment of risks before and after mitigation were not clearly defined.
 

	Investigation By:
Person responsible 
	Derek Lamb

	Investigation Issue id
(if applicable)
Root Cause Analysis 
	We did not have a risk assessment in place for the service and production of the Tom Thumb. 
It should have been present in Intrastats in the document Index. 
The Tom Thumb is a legacy product manufactured by Viamed. The product itself had an overall risk assessment in the system, but it did not cover service and production. 


	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action Issue ID (if applicable):
Relevant and Proportionate Corrective Action 
	Issue ID 155069
Issue ID 155535
Issue ID 155533
Issue ID 155069
Issue ID 154979
Issue ID 154721
Issue ID 154720


	Time Scale for Corrective Action
Time for completion of all identified actions 
	Immediate Action to be completed by 19th October 2019
Corrective action to be completed by 19th October 2019

	Corrective Action:
	Immediate Action
We will write an up to date Risk Assessment for all stages of the production, service and QA of the Tom Thumb. Which will be added to the product Risk Assessment in the system. 
We no longer produce this item, but for completeness we will make the Risk Assessment cover all areas to ensure nothing is missed. 
Corrective Action
We will check all production, servicing and QA areas / products for an up to date Risk Assessment, verifying they are fit for purpose and easily accessible to staff. 
Then add these to the Required Reading to make sure they have been read and that there is a record of it in Intrastats. 


	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


