    Oxygen Monitor Software specification




Function
Comments
Y/N
Action Required

1

Automatic Calibration

N
None Required


a
One & Two point calibration Air & 100% Oxygen

Y
Change membrane symbol Bell to Cal


b
Two push for calibration to prevent accidental calibration  in use
This can be two step i.e menu; calibrate Y/N
Y
Sometimes unit reads Calibration error but after return reads calibration value correctly

Span cal not working correctly.

Should Call 100 then Cal 21 then reads out span ratio.



The actual menu and function are OK 


c
Detects the standing sensor output and determines whether calibration is in Air or  100% 02
   Under 13mV cal in Air. Over  45mV cal in 100% oxygen
N
At 14mV shows Cal error but then calibrates to 21% Calibrates over 50mV


d
Detects the output i.e. l0.5mV +/-  3mV for Air & 50 mV +/-   l5mV for 100% 02
Under 9mV in Air or 45 mV in 100% “ check sensor”
N
Calibrates down very low levels on both 21 & 100% calibrations


e
In two point calibration it should be able to calculate the span ratio
 Reading in 100% O2 / reading in Air. Span should

 be 4.785 +/- 3%
N
How does this function work


f
Calibration should only disable alarms for the time taken to calibrate

Y




If we use the Audio silence as a Calibration control how do we silence Alarms   See g for a solution We could then move cal to Menu 


g
On switch on the monitor should ask the user if calibration is required. (It may be sitting in enriched atmosphere
If sensor output is over 12mV it should repeat

 message 
N



h
Calibration due if not automatically detected i.e every 8 hours
Using internal clock
N



i
Memory of initial output
Only if sensor is chipped
N



j
On the first calibration the sensor remembers the sensor output in Air and the sensor output in 100% 02
Only if sensor is chipped
N



k
If this is built into the sensor it can be done during manufacture
Only if sensor is chipped
N



l
Memory of use.
Only if sensor is chipped
N



m
It should be possible to add manufacturing date and then calculate the actual % hours of 02 use
Only if sensor is chipped




n
Serial number

The sensor if it has a chip should have the serial number and  manufacturing date and use in 02 memorised
Only if sensor is chipped
N


2

Alarms





a
1c & 1d can be overcome if on switch on the Auto Alarm is activated and only deactivated on Calibration 


b
High Alarm

Y



c
Settable 0-100% in 1% steps 

Y



d
Should not be able to set High alarm below the Low alarm setting

Y



e
High alarm set above 120% should switch off High Alarm
In this position the alarm should be silent but should have a visual Red flashing alarm digits.




f
It is possible to switch off Hi Alarm; Only small notification Off on Hi Alarm position. Instrument will nor read above 100%. Instrument must register above 100%. e.g. CPAP using 100%


g





3

Low Alarm





a
Settable 0-100% in 1% steps 

Y



b
Should not be able to set Low alarm above the High alarm  setting

Y



c
Safety alarm





d
This should be set at 18% regardless of Low alarm set  
Lo Alarm can be turned off
N



e
It should be impossible to disable Audio or visual alarm indication with this alarm.

N



f
A method of disabling this feature is required in service section for non medical  analysers. Message should be on screen  “Sensor alarm disabled”

Y
Needs message Sensor alarm dissabled

4

Indicators





a
Low battery warning

Y



b
Battery remaining life

Y



c
On Alarm Red LED or Red flashing alarm digits
Digits and alarm levels go Red
Y



d
Flashing Alarm indicator digits High and  Low
Digits Flash Red
Y




Indication of Alarm off

Y




Indication of alarm silenced.

Y


5
a
Alarm silence





b
1 push 60 secs

N
Needs to be added 


c
2 push 90 secs

N
Needs to be added 


d
3 push 120 secs maximum

N
Needs to be added 


e
Use shrinking bar if possible or digital countdown.

N


6

Output





a
USB

Y



b
Analogue 0-1 volt DC




7

Display





a
Digital Large Digits with alarm settings in small digits

Y



b
Analogue.

N



b
 A semi circle clock with pointer. Between alarm set points in Green
Not checked
N


8

Data management





a
Stores readings every 30 seconds with alarm settings recorded
At present 4 secs. 30 secs is OK
Y



b
Show graph of 02 readings

Y



c
Shows alarm limits 

Y



d
Shows calibration points with Date & Time
Date function only if unit has a calender
N



e
Depending on memory available it should be continuous with first on being replaced by last reading.





f
No need for patient identity

N


9

Physical





a
Three pole jack plug 

Y
Seems easily dislodged


b
Should work on sensors with two pole jack socket

Y



c
Remote sensor can incorporate a chip . Depending on the memory it could  hold much of the data  requested in above.

N


10

Battery Life.
100hrs




a
Can be increased using re-chargeable batteries

N



b
Requires external charger or charger cradle

N



c
Table top version could use D cells and/or re-chargeable

N


11

Back light;


Not required in 1st moderl


a
ON/OFF

N
Not required in 1st moderl


b
Timed after a key press

N
Not required in 1st model

12

Power supply

N
Not required in 1st model


a
Internal

N
Not required in 1st model


b
External

N
Not required in 1st model

13

Pressure compensation

N
Not required in 1st model


a
Sensor with built in pressure transducer
Electronics compensate for deviations from 1Bar
N
Not required in 1st model


b
Needs a gas passage from front to rear

N
Not required in 1st model


c
Requires multi pin connector

N
Not required in 1st model

14

Sensor response time

N
Not required in 1st model


a
Can be varied with thickness of membrane

N
Not required in 1st model






















OXYGEN MONITOR SOFTWARE080722.doc              07/21/08                                            6

