Hi Victoria

You have requested post market surveillance data for recent purchase orders: POR12349 through to POR12444.

The following are the replies to the posed questions:

- Did you receive your goods in satisfactory condition?

Yes, this has improved over the past few months. Quality of packaging and collation of paperwork correlated to packaged stock has improved.

See end column in the table below for goods-in QA failures. This procedure is to verify the satisfactory functional condition of the sensors and to ensure that they meet the output specifications.

- Do you know of any possible relevance to safety regarding our medical devices? No

- Have you had any issues relating to clinical safety or performance (i.e. residual risks) of our medical devices in accordance with its approved labelling? No, none reported.

As far as the orders are concerned the confirmation response is very good and the delivery schedules have improved significantly.

|  | Purchase order | Qty | Product | Order date | Confirmed | Confirmed response | Scheduled delivery date | Delivery date | Receipt | Days to deliver | QA failures |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Medical | POR12349 | 100 | UFO 130-2 | 10/07/19 | 11/07/19 | +1 day | 30/08/19 | 27/08/19 | 3 days early | 46 days | 0 |
| Medical | POR12359 | 25 | AX300-I | 15/07/19 | 15/07/19 | Same day | 30/08/19 | 28/08/19 | 2 days early | 44 days | 1 |
| Medical | POR12359 | 50 | MX300-I | 15/07/19 | 15/07/19 | Same day | 30/08/19 | 24/09/19 | 25 days overdue | 40 days | 0 |
| Medical | POR12360 | 100 | R17MED | 15/07/19 | 15/07/19 | Same day | 30/08/19 | 27/08/19 | 3 days early | 43 days | 0 |
| Medical | POR12360 | 25 | T1 | 15/07/19 | 15/07/19 | Same day | 30/08/19 | 06/08/19 | 24 days early | 22 days | 0 |
| Medical | POR12360 | 100 | UFO 130-2 | 15/07/19 | 15/07/19 | Same day | 30/08/19 | 27/08/19 | 3 days early | 43 days | 0 |
| Medical | POR12361 | 200 | R-17MED | 15/07/19 | 15/07/19 | Same day | 06/09/19 | 03/09/19 | 3 days early | 51 days | 0 |
| Automotive | POR12362 | 50 | R-17A | 15/07/19 | 15/07/19 | Same day | 06/09/19 | 20/08/19 | 17 days early | 37 days | 1 |
| Automotive | POR12362 | 100 | R-17AV | 15/07/19 | 15/07/19 | Same day | 06/09/19 | 28/08/19 | 9 days early | 45 days | 0 |
| Medical | POR12369 | 100 | UFO 130-2 | 22/07/19 | 22/07/19 | Same day | 06/09/19 | 27/08/19 | 10 days early | 37 days | 0 |
| Medical | POR12381 | 100 | UFO 130-2 | 01/08/19 | 01/08/19 | Same day | 20/09/19 | 23/09/19 | 3 days overdue | 54 days | 0 |
| Medical | POR12383 | 60 | R-36 | 01/08/19 | 05/08/19 | +4 days | 27/09/19 | 21/08/19 | 37 days early | 17 days | 0 |
| Medical | POR12383 | 50 | UF0 130-2 | 01/08/19 | 05/08/19 | +4 days | 27/09/19 | 02/10/19 | 5 days overdue | 59 days | 0 |
| Automotive | POR12415 | 300 | R-22A | 22/08/19 | 26/08/19 | +4 days | 11/10/19 | 09/10/19 | 2 days early | 45 days | 0 |
| Automotive | POR12415 | 300 | R-22AV | 22/08/19 | 26/08/19 | +4 days | 11/10/19 | 09/10/19 | 2 days early | 45 days | 0 |
| Medical | POR12423 | 100 | UF0 130-2 | 02/09/19 | 03/09/19 | +1 day | 25/10/19 | 09/10/19 | 16 days early | 53 days | 0 |
| Automotive | POR12432 | 50 | R-17AH | 05/09/19 | 05/09/19 | Same day | 25/10/19 | Not received |  |  |  |
| Medical | POR12444 | 100 | UFO 130-2 | 16/09/19 | 24/09/19 | +8 days | 08/11/19 | Not received |  |  |  |
| Automotive | POR12444 | 50 | R-17AH | 16/09/19 | 24/09/19 | +8 days | 08/11/19 | Not received |  |  |  |

As the orders are recent orders, there hasn’t been a sufficient period for the products to have been in use at all or for a significant period. I would suggest that data be analyzed once the goods have been in use for 12 months. However, the following global historical failure rates may be of interest.

|  | Failure during QA inspection | Failure in use |
| --- | --- | --- |
| AX300-I | 0.86% | 2.17% |
| MX300-I | 1.6% | 3.80% |
| UFO 130-2 | 0.0% | 0.2% |
| R-17MED | 0.66% | 0.34% |
| R-36MED | 0.33% | 0.07% |
| T-1 | 0.49% | 0.16% |
| R-17A | 0.87% | 0.45% |
| R-17AV | 2.0% | 0.12% |
| R-22A | 0.8% | 1.22% |
| R-17AH | 5.28% | 3.85% |
| R-22AV | 1.07% | 0.83% |

NOTES

Failure out of use includes out of warranty figures.

The stats for automotive sensors are skewed. Over the recent years we have had high failure rates, but these have been masked by the historical good performance.