- Certificate of Calibration

- Certificate number:

- Model Number: AII-2000

Part Number:

Serial number:

This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with procedure: VM3COP??.??.

| Equipment used: | Time Electronics LTD Microcal 1030 – Calibrated Voltage source (CE076) - calibration traceable to UKAS calibration standards. |
| --- | --- |

| Pass: Observed Value within Expected Value Fail: Observed Value outside of Expected Value | Expected Value (± 2%) | Pass / Fail |
| --- | --- | --- |
| Unit passes internal start-up tests | PASS |  |
| Reading in air after air calibration (Calibration) | 20.9% oxygen |  |
| Reading after exposure to 100% oxygen (Linearity) | 100% oxygen |  |
| Reading in air after exposure to 100% oxygen (Accuracy) | 20.9% oxygen |  |
| Reading in 100% oxygen after 100% oxygen calibration (Calibration) | 100% |  |
| Reading fluctuation on 100% oxygen for 5 minutes (Noise) | < 1.0% oxygen |  |
| Reading over 5 minutes on 100% oxygen (Drift) | 100% oxygen |  |
| Reading in air after 100% oxygen calibration (Linearity) | 20.9% |  |
| Reading in 100% oxygen after exposure in air (Accuracy) | 100% oxygen |  |
| Overall inspection for physical defects |  |  |

Name:

Date:

Signed: 	______________________