- Certificate of Calibration

- Certificate number:

- V1000 Foetal Heart Simulator

Part Number: 1410000

Serial number:

This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with procedure: VM3COP55.01.

| Equipment used: | Sonicade FM6 foetal monitor (CE051) |
| --- | --- |
|  | TTi TF930 3GHz counter (CE185) - calibration traceable to UKAS calibration standards. |

| V1000 BPM setting | - Period values stated in milliseconds |  |  |  | Result PASS/FAIL |
| --- | --- | --- | --- | --- | --- |
|  | Target | Lower limit | Upper limit | Actual |  |
| 60 | 1000.00 | 990.00 | 1010.00 |  | PASS |
| 90 | 666.70 | 660.03 | 673.37 |  | PASS |
| 120 | 500.00 | 495.00 | 505.00 |  | PASS |
| 150 | 400.00 | 396.00 | 404.00 |  | PASS |
| 180 | 333.30 | 329.97 | 336.63 |  | PASS |
| 210 | 285.70 | 282.84 | 288.25 |  | PASS |

Name:

Date:

Signed: 	______________________