Post Market Surveillance Review to 31st December 2020

4310001 Flowsensor SpiroTrue A. Box of 5

Bluepoint Medical ref.: 1030132006

First sold: 12th January 2009

General sales suspended 18th December 2013 due to MHRA stance on single use designation. Reinstated when covered by Bluepoint Medical CE certification.

Total sales to date: 4,686 boxes, 111 different customers across 34 countries

Total failures: 6 boxes

Total returns: 6 boxes

Total customer failures: 0.128%

Liverpool Heart & Chest Hospital SRS63066, received 1 box 10th October 2011

Failed calibration. Also reported sensors generally have shorter life than Draeger OEM sensors.

Result: Customer lost, approximately 100 boxes per year. Resumed using Draeger OEM sensors, will try and regain the sales with Flowsensor SpiroTrue A.

Dorset County Hospital SRS63869, received 1 box 3rd June 2013

Failed calibration.

Result: Customer lost, approximately 12 boxes per year. Will try and regain the sales with Flowsensor SpiroTrue A.

Papworth Hospital SRS64002, received 4 boxes 11th September 2013

Failed calibration when used on Evita 4s.

Result: Customer lost, approximately 120 boxes per year. They have tried both SpiroTrue A and SpiroTrue A (PC). Will try and regain the sales with Flowsensor SpiroTrue A.

For the period from 12th September 2013 to 31st December 2020, there has not been any new reported failures or failures that refer to the claims listed above.

The design, functionality and reliability of the sensor is very good. We do include feedback forms within each box of sensors for end users to be able to easily report possible issues and/or for confirming compatibility. Unfortunately, we get very few feedback forms returned as the customers are happy with the sensors and have nothing to report.

Applications:

| Manufacturer | Description |  |
| --- | --- | --- |
| Acutronic | 1601+ | Flow Sensor |
| Acutronic | SpiroQuant A | Flow Sensor |
| Asmuth GmbH Medinzintechnik | AS432+ | Flow Sensor |
| Asmuth GmbH Medinzintechnik | AS8403735 | Flow Sensor |
| Asmuth GmbH Medinzintechnik | SpiroTrue A | Flow Sensor |
| Biometric Cables | 910246001 | Flow Sensor |
| Biometric Cables | FS1 | Flow Sensor |
| Bluepoint Medical | 1030132006 | Flow Sensor |
| Bluepoint Medical | 6871980 (NOTE Viamed does not have RFID) | Flow Sensor |
| Bluepoint Medical | PM 8030 | Monitor |
| Bluepoint Medical | SpiroTrue A | Flow Sensor |
| Draeger | Spirolog 84 03 735 | Flow Sensor |
| Draeger | 8403735 | Flow Sensor |
| Draeger | Anemone | Monitor |
| Draeger | Apollo | Anaesthesia System |
| Draeger | AV1 | Anaesthesia System |
| Draeger | Cato | Anaesthesia System |
| Draeger | Cicero | Anaesthesia System |
| Draeger | D-9118-2009 | Flow Sensor |
| Draeger | Dura | Ventilator |
| Draeger | Eva | Ventilator |
| Draeger | Evita | Ventilator |
| Draeger | Evita 4 | Ventilator |
| Draeger | Evita XL | Ventilator |
| Draeger | Fabius GS | Anaesthesia System |
| Draeger | Infinity ID (NOTE Viamed ver. does not have RFID) | Flow Sensor |
| Draeger | Infinity V500 (NOTE Viamed ver. does not have RFID) | Anaesthesia System |
| Draeger | Julian | Anaesthesia System |
| Draeger | MT-2086-2007 (NOTE Viamed ver. does not have RFID) | Flow Sensor |
| Draeger | PM 8030 | Monitor |
| Draeger | PM 8035 | Monitor |
| Draeger | PM 8050 | Monitor |
| Draeger | PM 8060 | Monitor |
| Draeger | Primus | Anaesthesia System |
| Draeger | Savina | Ventilator |
| Draeger | Spirolog | Flow Sensor |
| Draeger | Spirolog 8403735 | Flow Sensor |
| Draeger | Sulla | Anaesthesia System |
| Envitec | 1000470 | Flow Sensor |
| Envitec | SpiroQuant A+ | Flow Sensor |
| Integral Process | 0032 00 02 | Flow Sensor |
| Integral Process | 65025 | Flow Sensor |
| Integral Process | 8403735 | Flow Sensor |
| Intersurgical | 2777000 | Flow Sensor |
| Intersurgical | e-FlowSensor | Flow Sensor |
| Lub Dub Medical | OX-2529 | Flow Sensor |
| Medacx | M-45-00-0005 | Flow Sensor |
| Medacx | SpiroQuant A | Flow Sensor |
| Proact Medical | 1000470 | Flow Sensor |
| Proact Medical | SpiroQuant A+ | Flow Sensor |
| Viamed | 1030132006 | Flow Sensor |
| Viamed | Flowsensor A | Flow Sensor |
| Viamed | SpiroTrue A | Flow Sensor |

Reviewed by: Steve Nixon

Director – Viamed Ltd.

Date: 	          5th February 2021