Post Market Surveillance Review to 31st December 2020

4310003 Flowsensor SpiroTrue H. Box of 6

Bluepoint Medical ref.: 103013200

First sold: 8th May 2009

Total sales to date: 6,790 boxes, 49 different customers across 29 countries

Total failures: 2 boxes

Total returns: 2 packs – failed calibration on Hamilton G5 ventilator, reported then returned to Viamed for credit 10/12/2009. Sensors supplied to Bluepoint Medical for evaluation 29/12/2009

Customer: Royal Sussex County Hospital    SRS61983

Total customer failures: 0.03%

For the period from 30th December 2009 to 31st December 2020, there has not been any new reported failures or failures that refer to the claims listed above.

The design, functionality and reliability of the sensor is very good. We do include feedback forms within each box of sensors for end users to be able to easily report possible issues and/or for confirming compatibility. Unfortunately, we get very few feedback forms returned as the customers are happy with the sensors and have nothing to report.

Applications:

| Manufacturer | Description |  |
| --- | --- | --- |
| Biometric Cables | 910246002 | Flow Sensor |
| Biometric Cables | FS2 | Flow Sensor |
| Draeger | Oxilog 2000 Plus (if disposable circuits used) | Ventilator |
| Draeger | Oxilog 3000 (if disposable circuits used) | Ventilator |
| Draeger | Oxilog 3000 Plus (if disposable circuits used) | Ventilator |
| Draeger | VentStar | Ventilator |
| Hamilton Medical | 155362 | Flow Sensor |
| Hamilton Medical | 279331 | Flow Sensor |
| Hamilton Medical | 53120 | Flow Sensor |
| Hamilton Medical | Amadeus | Ventilator |
| Hamilton Medical | C1 | Ventilator |
| Hamilton Medical | C2 | Ventilator |
| Hamilton Medical | G5 | Ventilator |
| Hamilton Medical | Galileo | Ventilator |
| Hamilton Medical | H72 | Ventilator |
| Hamilton Medical | H7200 | Ventilator |
| Hamilton Medical | H7400 AccuTach | Flow Sensor |
| Hamilton Medical | MR1 | Ventilator |
| Hamilton Medical | Raphael | Ventilator |
| Hamilton Medical | S1 | Ventilator |
| Hamilton Medical | T1 | Ventilator |
| Hamilton Medical | Veolar | Ventilator |
| Hamilton Medical | VR-1 | Ventilator |
| Lub Dub Medical | OX-2530 | Flow Sensor |
| Med-Shopping | MAC-VTA-OTH-FSH7200 | Flow Sensor |
| Viamed | 3030131005 | Flow Sensor |
| Viamed | 1030132000 | Flow Sensor |
| Viamed | Flowsensor H | Flow Sensor |
| Viamed | SpiroTrue H | Flow Sensor |

Reviewed by:

Steve Nixon

Director – Viamed Ltd.

Date:               5th February 2021