## Sheet1
SENSOR EVALUATION TABLE
ID	CELL CLASS	CUSTOMER	CONTROL #	DATE RECEIVED	DATE EVALUATED	SERIAL #	QTY	DATE SHIPPED	VISUAL INSPECTION	DEFECT CODE	TEST DATA	DISPOSITION CODE	ANALYSIS NOTES	Field1
25435	R17MED	Viamed	31403	03/05/2019	03/07/2019	See hard copy	8	04/13/18	Good	O5	See hard copy	RUW	Two sensors passed all test and 6 failed test for linearity error more than 1.5%
25436	R22MED	Viamed	31403	03/05/2019	03/07/2019	137413	1	04/20/18	Good	O5	See hard copy	RUW	Sensor failed for linearity error more than 1.5%. See test data.
25438	R17AH	Viamed	30834	02/19/19	03/08/2019	See hard copy	102	See hard copy	Good	O7	See hard copy	RUW	There were 60 sensors that are within TAI test limit but borderline high and 42pcs above test limit.
25439	R22AHJR	Viamed	30834	02/19/19	03/08/2019	See hard copy	286	See hard copy	Good	O7	See hard copy	RUW	There were 66 sensors are within TAI test limit but borderline high and 220pcs. Above TAI test limit.
25450	R17MED	Viamed	31757	04/01/2019	04/08/2019	See hard copy	18	See hard copy	Good	O1	See hard copy	RUW	Two sensors passed all test, nine with zero output, three high output, and four failed test for non-linearity.
25451	R22MED	Viamed	31757	04/01/2019	04/08/2019	221,812,232,369	2	See hard copy	Good	O1	See hard copy	RUW	One sensor had no output in air and second failed for linearity error.
25452	R22A	Viamed	31748	04/01/2019	04/09/2019	See hard copy	122	See hard copy	Good	O3	See hard copy	RUW	High = 50pcs. Unstable 40pcs. Non-linear 1pc. Vacuum leak 3pcs. NIF 28pcs.
25453	R22AV	Viamed	31748	04/01/2019	04/09/2019	See hard copy	29	See hard copy	Good	O3	See hard copy	RUW	High = 15pcs. Vacuum leak = 4pcs. Unstable = 8pcs. Dead = 2pcs.
25454	R17ALV	Viamed	31748	04/01/2019	04/09/2019	See hard copy	7	02/02/2019	Good	O3	See hard copy	RUW	High output = 6pcs. Unstable = 1
25455	R17AH	Viamed	31748	04/01/2019	04/09/2019	242001	1	11/30/18	Good	O3	14mv	RUW	Sensor output in air is above test limit.
25468	R22AH	Viamed	33217	05/28/19	05/29/19	See hard copy	48	See hard copy	Good	O3	See hard copy	RUW	Thirthy three sensors passed all test but 1mv or less away from maximum test limit. Fifteen failed for high unstable output and non linearity.
25469	R21AHAIK	Viamed	33217	05/28/19	05/29/19	See hard copy	5	06/29/18	Good	O3	See hard copy	RUW	All five sensors output in air are above test limit.