## Summary
`0110017 | Oxygen sensor R-17MED | 166
`0110022 | Oxygen sensor R-22MED | 51
`0110024 | Oxygen sensor R-24MED | 40
`0110026 | Oxygen sensor R-26MED | 1
`0110051 | Oxygen sensor T-1 | 3
`0110057 | Oxygen sensor T-7 | 5
`0110081 | Oxygen sensor UFO 130-2 | 21
| Total | 287

## R-17MED
|  |  |  | Description | Serial number | Fault
546 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572182 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
549 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572152 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
565 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572117 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
552 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572111 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
563 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572105 | Output ( 7mV - 13mV ) 14.4 Failed during QA, did not reach end user
547 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572104 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
545 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572092 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
550 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572072 | Output ( 7mV - 13mV ) 14.4 Failed during QA, did not reach end user
561 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572056 | Output ( 7mV - 13mV ) 14.0 Failed during QA, did not reach end user
560 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572031 | Output ( 7mV - 13mV ) 13.3 Failed during QA, did not reach end user
553 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572024 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
554 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 572004 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
559 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571879 | Output ( 7mV - 13mV ) 14.0 Failed during QA, did not reach end user
564 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571536 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
555 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571534 | Output ( 7mV - 13mV ) 14.5 Failed during QA, did not reach end user
551 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571533 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
557 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571517 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
556 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571514 | Output ( 7mV - 13mV ) 14.1 Failed during QA, did not reach end user
548 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 571363 | Output ( 7mV - 13mV ) 14.9 Failed during QA, did not reach end user
562 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 566507 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
558 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 566481 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
445 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 566162 | Output ( 7mV - 13mV ) 14.6 Failed during QA, did not reach end user
447 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 566154 | Output ( 7mV - 13mV ) 13.4 Failed during QA, did not reach end user
442 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 566017 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
446 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565879 | Output ( 7mV - 13mV ) 13.4 Failed during QA, did not reach end user
444 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565867 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
443 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565744 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
451 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565712 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
566 | BOX765 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565706 | Output ( 7mV - 13mV ) 14.2 Failed during QA, did not reach end user
449 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565698 | Output ( 7mV - 13mV ) 14.3 Failed during QA, did not reach end user
450 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565626 | Output ( 7mV - 13mV ) 14.1 Failed during QA, did not reach end user
448 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 565625 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
452 | BOX755 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 554104 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
386 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 524630 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
245 | BOX745 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 523016 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
247 | BOX745 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 523014 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
246 | BOX745 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 523011 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
244 | BOX745 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 522810 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
148 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 491503 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
143 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 479839 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
387 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 479763 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
152 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 477702 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
142 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 475921 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
139 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 475658 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
145 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 475435 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
150 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 475116 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
388 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 471436 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
392 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 470275 | Connector Fault
411 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 470063 | Connector Fault
390 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 470012 | Connector Fault
389 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 469813 | Non Linear
407 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 469717 | Connector Fault
151 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 466389 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
391 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 466052 | Connector Fault
417 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465878 | Non Linear
404 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465833 | Connector Fault
410 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465687 | Connector Fault
396 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465646 | Connector Fault
419 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465626 | Connector Fault
154 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465558 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
161 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465350 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
144 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465333 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
416 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465318 | Connector Fault
149 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 465296 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
398 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 453182 | Connector Fault
393 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 453109 | Connector Fault
156 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 453094 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
415 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 453064 | Connector Fault
394 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 453017 | Connector Fault
409 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 453005 | Connector Fault
162 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452947 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
405 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452846 | Non Linear
412 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452756 | Non Linear
403 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452718 | Connector Fault
153 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452651 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
157 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452629 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
413 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452472 | Connector Fault
397 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452325 | Connector Fault
155 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 452272 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
400 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 451561 | Connector Fault
160 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 451292 | Output ( 7mV - 13mV ) Intermittent output. Failed during QA, did not reach end user
418 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 451286 | Connector Fault
395 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 450900 | Connector Fault
426 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 419947 | Non Linear
430 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 419893 | Non Linear
141 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 419806 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
140 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417193 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
429 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417189 | Connector Fault
427 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417183 | Non Linear
146 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417096 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
147 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417060 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
420 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417013 | Connector Fault
408 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417004 | Connector Fault
433 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 417002 | Non Linear
425 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416977 | Non Linear
428 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416976 | Unstable
402 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416931 | Unstable
406 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416898 | Non Linear
421 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416873 | Non Linear
401 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416866 | Unstable
414 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416849 | Connector Fault
424 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416843 | Non Linear
399 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416842 | Non Linear
422 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416838 | Non Linear
434 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416791 | Connector Fault
431 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416712 | Non Linear
159 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416697 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
158 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416675 | Output ( 7mV - 13mV ) 13.3 Failed during QA, did not reach end user
432 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416613 | Connector Fault
423 | BOX750 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 416595 | High Output
110 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411413 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
108 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411409 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
103 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411345 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
100 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411268 | Output ( 7mV - 13mV ) 13.3 Failed during QA, did not reach end user
102 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411168 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
101 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411139 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
105 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411025 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
104 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 411008 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
115 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410943 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
114 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410894 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
113 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410860 | Output ( 7mV - 13mV ) 13.3 Failed during QA, did not reach end user
111 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410748 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
116 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410740 | Output ( 7mV - 13mV ) 31.7 Failed during QA, did not reach end user
106 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410510 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
107 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 410026 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
117 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 409922 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
120 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 409883 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
119 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 409847 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
112 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 409830 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
109 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 409824 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
118 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 409797 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
74 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 404513 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
75 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397212 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
76 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397194 | Output ( 7mV - 13mV ) 15.0 Failed during QA, did not reach end user
79 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397193 | Output ( 7mV - 13mV ) 14.1 Failed during QA, did not reach end user
80 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397183 | Output ( 7mV - 13mV ) 14.9 Failed during QA, did not reach end user
77 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397182 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
78 | BOX718 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397167 | Output ( 7mV - 13mV ) Erratic output. Failed during QA, did not reach end user
54 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397110 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
64 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397109 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
55 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397068 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
59 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397041 | Output ( 7mV - 13mV ) 14.3 Failed during QA, did not reach end user
60 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 397033 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
57 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 396943 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
56 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 396879 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
58 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 390016 | Output ( 7mV - 13mV ) 14.2 Failed during QA, did not reach end user
61 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 390004 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
63 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 389997 | Output ( 7mV - 13mV ) 14.2 Failed during QA, did not reach end user
62 | BOX715 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 389993 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
43 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386917 | Output ( 7mV - 13mV ) 13.8 Failed during QA, did not reach end user
41 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386883 | Output ( 7mV - 13mV ) 13.1 Failed during QA, did not reach end user
42 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386828 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
48 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386632 | Output ( 7mV - 13mV ) 14.3 Failed during QA, did not reach end user
50 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386631 | Output ( 7mV - 13mV ) 10.9 Failed during QA, did not reach end user
46 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386621 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
51 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386613 | Output ( 7mV - 13mV ) 14.0 Failed during QA, did not reach end user
52 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386608 | Output ( 7mV - 13mV ) 14.2 Failed during QA, did not reach end user
47 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386607 | Output ( 7mV - 13mV ) 14.5 Failed during QA, did not reach end user
53 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386603 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
49 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386597 | Output ( 7mV - 13mV ) 14.1 Failed during QA, did not reach end user
44 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386595 | Output ( 7mV - 13mV ) 14.3 Failed during QA, did not reach end user
39 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386573 | Output ( 7mV - 13mV ) No output. Failed in Q.A. n Did not reach end user
40 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 386565 | Output ( 7mV - 13mV ) 14.3 Failed during QA, did not reach end user
163 | BOX737 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 330885 | Non-linear
45 | BOX707 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 315084 | Fluctuating output during use
99 | BOX727 | Not Yet Shipped | 0110017 | Oxygen sensor R-17MED | 305104 | Connector Fault

## R-22MED
|  |  |  | Description | Serial number | Fault
574 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 589425 | Output ( 7mV - 13mV ) 14.4 Failed during QA, did not reach end user
571 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 589346 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
579 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 589271 | Output ( 7mV - 13mV ) Erratic. Failed during QA, did not reach end user
573 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582577 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
572 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582575 | Output ( 7mV - 13mV ) 13.8 Failed during QA, did not reach end user
568 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582565 | Output ( 7mV - 13mV ) 13.3 Failed during QA, did not reach end user
569 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582548 | Output ( 7mV - 13mV ) 14.2 Failed during QA, did not reach end user
577 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582546 | Output ( 7mV - 13mV ) 13.8 Failed during QA, did not reach end user
567 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582540 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
575 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582536 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
578 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582532 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
570 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582530 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
576 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 582529 | Output ( 7mV - 13mV ) 14.6 Failed during QA, did not reach end user
453 | BOX755 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 540641 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
455 | BOX755 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 532179 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
456 | BOX755 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 532165 | Output ( 7mV - 13mV ) 14.2 Failed during QA, did not reach end user
454 | BOX755 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 532164 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
249 | BOX745 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 531512 | Output ( 7mV - 13mV ) 13.8 Failed during QA, did not reach end user
167 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 522182 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
250 | BOX745 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 520470 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
248 | BOX745 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 519905 | Output ( 7mV - 13mV ) 14.5 Failed during QA, did not reach end user
170 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 505405 | Output ( 7mV - 13mV ) 13.9 Failed during QA, did not reach end user
171 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 505368 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
168 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 505344 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
169 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 505343 | Output ( 7mV - 13mV ) 13.4 Failed during QA, did not reach end user
172 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 483893 | Output ( 7mV - 13mV ) Intermittent output.
438 | BOX750 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 468048 | Output ( 7mV - 13mV ) Non-linear. Failed during QA, did not reach end user
439 | BOX750 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 459994 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
580 | BOX765 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 457482 | Output ( 7mV - 13mV ) Non-linear. Failed during QA, did not reach end user
173 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 456137 | Output ( 7mV - 13mV ) 2.7 Failed during QA, did not reach end user
437 | BOX750 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 447988 | Output ( 7mV - 13mV ) Non-linear. Failed during QA, did not reach end user
436 | BOX750 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 447922 | Output ( 7mV - 13mV ) Non-linear. Failed during QA, did not reach end user
435 | BOX750 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 442577 | Output ( 7mV - 13mV ) Non-linear. Failed during QA, did not reach end user
164 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 426140 | Output ( 7mV - 13mV ) No output. Failed during QA, did not reach end user
165 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 418832 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
166 | BOX737 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 418597 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
123 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406295 | Output ( 7mV - 13mV ) 14.6 Failed during QA, did not reach end user
130 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406292 | Output ( 7mV - 13mV ) 14.3 Failed during QA, did not reach end user
128 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406283 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
124 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406279 | Output ( 7mV - 13mV ) 13.4 Failed during QA, did not reach end user
127 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406276 | Output ( 7mV - 13mV ) 13.2 Failed during QA, did not reach end user
129 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406267 | Output ( 7mV - 13mV ) 15.5 Failed during QA, did not reach end user
121 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406249 | Output ( 7mV - 13mV ) 13.3 Failed during QA, did not reach end user
126 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406232 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
122 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406225 | Output ( 7mV - 13mV ) 13.8 Failed during QA, did not reach end user
125 | BOX727 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406224 | Output ( 7mV - 13mV ) 13.5 Failed during QA, did not reach end user
85 | BOX722 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406217 | Output ( 7mV - 13mV ) 13.7 Failed during QA, did not reach end user
84 | BOX722 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406213 | Output ( 7mV - 13mV ) 13.4 Failed during QA, did not reach end user
83 | BOX722 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406210 | Output ( 7mV - 13mV ) 13.6 Failed during QA, did not reach end user
82 | BOX722 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 406206 | Output ( 7mV - 13mV ) 13.8 Failed during QA, did not reach end user
65 | BOX715 | Not Yet Shipped | 0110022 | Oxygen sensor R-22MED | 394703 | Defective connector, erratic output

## R-24
|  |  |  | Description | Serial number | Fault
463 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569492 | Faulty Sensor Cap Orientation
464 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569508 | Faulty Sensor Cap Orientation
475 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569512 | Faulty Sensor Cap Orientation
467 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569513 | Faulty Sensor Cap Orientation
469 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569516 | Faulty Sensor Cap Orientation
470 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569518 | Faulty Sensor Cap Orientation
471 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569527 | Faulty Sensor Cap Orientation
472 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569528 | Faulty Sensor Cap Orientation
473 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569534 | Faulty Sensor Cap Orientation
474 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569537 | Faulty Sensor Cap Orientation
466 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569539 | Faulty Sensor Cap Orientation
465 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569546 | Faulty Sensor Cap Orientation
476 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569548 | Faulty Sensor Cap Orientation
477 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569550 | Faulty Sensor Cap Orientation
478 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569552 | Faulty Sensor Cap Orientation
479 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569553 | Faulty Sensor Cap Orientation
480 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569554 | Faulty Sensor Cap Orientation
481 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569557 | Faulty Sensor Cap Orientation
482 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569568 | Faulty Sensor Cap Orientation
483 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569570 | Faulty Sensor Cap Orientation
484 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569572 | Faulty Sensor Cap Orientation
485 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569579 | Faulty Sensor Cap Orientation
486 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569584 | Faulty Sensor Cap Orientation
487 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569586 | Faulty Sensor Cap Orientation
488 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569587 | Faulty Sensor Cap Orientation
489 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569588 | Faulty Sensor Cap Orientation
490 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569589 | Faulty Sensor Cap Orientation
491 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569592 | Faulty Sensor Cap Orientation
492 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569594 | Faulty Sensor Cap Orientation
493 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569595 | Faulty Sensor Cap Orientation
494 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569598 | Faulty Sensor Cap Orientation
495 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569599 | Faulty Sensor Cap Orientation
496 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569601 | Faulty Sensor Cap Orientation
497 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569602 | Faulty Sensor Cap Orientation
498 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569603 | Faulty Sensor Cap Orientation
499 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569604 | Faulty Sensor Cap Orientation
500 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569611 | Faulty Sensor Cap Orientation
501 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569612 | Faulty Sensor Cap Orientation
502 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 569613 | Faulty Sensor Cap Orientation
468 | BOX758 | Not Yet Shipped | 0110024 | Oxygen sensor R-24MED | 571434 | Faulty Sensor Cap Orientation

## R-26MED
|  |  |  | Description | Serial number | Fault
251 | BOX745 | Not Yet Shipped | 0110026 | Oxygen sensor R-26MED | 527273 | No output. Failed during QA, did not reach end user

## T-1
|  |  |  | Description | Serial number | Fault
81 | BOX718 | Not Yet Shipped | 0110051 | Oxygen sensor T-1 | 329954 | Electrolyte leakage. Failed during QA, did not reach end user
174 | BOX737 | Not Yet Shipped | 0110051 | Oxygen sensor T-1 | 498617 | Electrolyte leakage. Failed during QA, did not reach end user
175 | BOX737 | Not Yet Shipped | 0110051 | Oxygen sensor T-1 | 498653 | No output. Failed during QA, did not reach end user

## T-7
|  |  |  | Description | Serial number | Fault
177 | BOX737 | Not Yet Shipped | 0110057 | Oxygen sensor T-7 | 427896 | No output. Failed during QA, did not reach end user
178 | BOX737 | Not Yet Shipped | 0110057 | Oxygen sensor T-7 | 427928 | No output. Failed during QA, did not reach end user
176 | BOX737 | Not Yet Shipped | 0110057 | Oxygen sensor T-7 | 498820 | High output. Failed during QA, did not reach end user
179 | BOX737 | Not Yet Shipped | 0110057 | Oxygen sensor T-7 | 498847 | High output. Failed during QA, did not reach end user
180 | BOX737 | Not Yet Shipped | 0110057 | Oxygen sensor T-7 | 498868 | High output. Failed during QA, did not reach end user

## UFO
|  |  |  | Description | Serial number | Fault
260 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 355460 | Erratic output
261 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 361216 | Erratic output
265 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 361227 | Erratic output
266 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 364173 | Erratic output
263 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 364232 | Erratic output
253 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372393 | Erratic output
255 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372413 | Erratic output
254 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372415 | Erratic output
259 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372424 | Erratic output
257 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372445 | Erratic output
258 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372456 | Erratic output
256 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372457 | Erratic output
264 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372459 | Erratic output
267 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372460 | Erratic output
262 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372462 | Erratic output
268 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372473 | Erratic output
252 | BOX745 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 372566 | Erratic output
457 | BOX755 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 479208 | Erratic output
440 | BOX750 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 500895 | Erratic output
581 | BOX765 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 541661 | Erratic output
441 | BOX750 | Not Yet Shipped | 0110081 | Oxygen sensor UFO 130-2 | 364241 | Erratic output